## BAM QC ST-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-22 |  | Model Number: | BAM-1020
TIME: | 12:20 |  | Serial Number: | P21275
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:24 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:24
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.85 | -0.8902% | 1.08%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 8.1 | 9.1 | -1
Certification Expiration: | 2015-01-21 | 8.2 | 8.8 | -0.6
|  | 8 | 9 | -1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 744 | 745 | -1 | 632
Certification Expiration: | 2015-01-21 | 744 | 745 | -1 | 41280
|  | 744 | 745 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-22 |  | Model Number: | BAM-1020
TIME: | 12:40 |  | Serial Number: | P22904
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:40 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:40
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.68 | 0.1199% | 0.05999%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 8.4 | 7.6 | 0.8
Certification Expiration: | 2015-01-21 | 8.5 | 7.6 | 0.9
|  | 8.4 | 7.5 | 0.9
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 744 | 745 | -1 | 632
Certification Expiration: | 2015-01-21 | 744 | 745 | -1 | 41280
|  | 744 | 745 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-22 |  | Model Number: | BAM-1020
TIME: | 12:05 |  | Serial Number: | P22809
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:06 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:06
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.98 | -1.649% | 1.86%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 8.1 | 8.3 | -0.2
Certification Expiration: | 2015-01-21 | 8.1 | 8.6 | -0.5
|  | 8.2 | 8.6 | -0.4
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 744 | 745 | -1 | 632
Certification Expiration: | 2015-01-21 | 744 | 745 | -1 | 41280
|  | 744 | 745 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-N
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-N |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-22 |  | Model Number: | BAM-1020
TIME: | 11:15 |  | Serial Number: | P22906
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:17 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:17
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.5 | 1.212% | -1.02%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 7.5 | 7.6 | -0.1
Certification Expiration: | 2015-01-21 | 7.5 | 7.8 | -0.3
|  | 7.6 | 8 | -0.4
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 745 | 0 | 632
Certification Expiration: | 2015-01-21 | 745 | 745 | 0 | 41280
|  | 745 | 745 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-22 |  | Model Number: | BAM-1020
TIME: | 10:55 |  | Serial Number: | P22808
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:03 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:03
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.79 | -0.536% | 0.7199%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 7.1 | 7.6 | -0.5
Certification Expiration: | 2015-01-21 | 7.2 | 7.6 | -0.4
|  | 7.2 | 8 | -0.8
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 745 | 0 | 632
Certification Expiration: | 2015-01-21 | 745 | 745 | 0 | 41280
|  | 745 | 745 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM
DATE: | 2014-04-25 |  | Model Number: | BAM-1020
TIME: | 11:35 |  | Serial Number: | M4791
0.3
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:42 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:41
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.41 | 1.767% | -1.56%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18 | 16.7 | 1.3
Certification Expiration: | 2015-01-21 | 18 | 16.7 | 1.3
|  | 18 | 16.7 | 1.3
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 742 | -1 | 635
Certification Expiration: | 2015-01-21 | 741 | 742 | -1 | 41283
|  | 741 | 742 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: NT-SE monitor could not be accessed due to high water surrounding it. |  |  |  |  | Rev. 3 3/2014

## BAM QC NT-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-25 |  | Model Number: | 1020
TIME: | 12:00 |  | Serial Number: | M4793
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:04 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:03
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.56 | 0.8454% | -0.6599%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.8 | 17.7 | 0.1
Certification Expiration: | 2015-01-21 | 17.8 | 17.7 | 0.1
|  | 17.7 | 17.5 | 0.2
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 740 | 741 | -1 | 636
Certification Expiration: | 2015-01-21 | 740 | 741 | -1 | 41284
|  | 740 | 741 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-25 |  | Model Number: | 1020
TIME: | 11:05 |  | Serial Number: | P21274
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:15 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:15
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.4 | 16.14 | 1.611% | -3.179%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 16.8 | 16.1 | 0.7
Certification Expiration: | 2015-01-21 | 16.9 | 16.3 | 0.6
|  | 16.9 | 16.3 | 0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 740 | 742 | -2 | 635
Certification Expiration: | 2015-01-21 | 740 | 742 | -2 | 41283
|  | 740 | 742 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-04-25 |  | Model Number: | 1020
TIME: | 12:10 |  | Serial Number: | M4792
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:07 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:07
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.66 | 0.2401% | -0.05999%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18.4 | 18.2 | 0.2
Certification Expiration: | 2015-01-21 | 18.4 | 18.2 | 0.2
|  | 18.4 | 18.2 | 0.2
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 741 | 741 | 0 | 636
Certification Expiration: | 2015-01-21 | 741 | 741 | 0 | 41284
|  | 741 | 741 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014