## BAM QC ST-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-06 |  | Model Number: | BAM-1020
TIME: | 10:55 |  | Serial Number: | P21275
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:58 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:58
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.59 | 0.6631% | -0.4799%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 13.8 | 13.3 | 0.5
Certification Expiration: | 2015-01-21 | 13.7 | 13.1 | 0.6
|  | 13.8 | 13.3 | 0.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 746 | -1 | 631
Certification Expiration: | 2015-01-21 | 745 | 746 | -1 | 41279
|  | 745 | 746 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-06 |  | Model Number: | BAM-1020
TIME: | 10:40 |  | Serial Number: | P22904
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:45 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:44
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.87 | -1.008% | 1.2%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 11.4 | 12.3 | -0.9
Certification Expiration: | 2015-01-21 | 11.5 | 12.3 | -0.8
|  | 11.5 | 12.1 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 745 | 0 | 632
Certification Expiration: | 2015-01-21 | 745 | 745 | 0 | 41280
|  | 745 | 745 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-09 |  | Model Number: | BAM-1020
TIME: | 12:05 |  | Serial Number: | P22809
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:10 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:09
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.83 | -0.7724% | 0.9598%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 21.6 | 22.2 | -0.6
Certification Expiration: | 2015-01-21 | 21.6 | 22.2 | -0.6
|  | 21.6 | 22.1 | -0.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 737 | 1 | 640
Certification Expiration: | 2015-01-21 | 738 | 737 | 1 | 41288
|  | 738 | 737 | 1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-N
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-N |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-06 |  | Model Number: | BAM-1020
TIME: | 10:30 |  | Serial Number: | P22906
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:27 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:28
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.52 | 1.09% | -0.8998%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 11.1 | 11.6 | -0.5
Certification Expiration: | 2015-01-21 | 11.1 | 11.6 | -0.5
|  | 11.1 | 11.6 | -0.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 746 | -1 | 631
Certification Expiration: | 2015-01-21 | 745 | 746 | -1 | 41279
|  | 745 | 746 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-06 |  | Model Number: | BAM-1020
TIME: | 10:15 |  | Serial Number: | P22808
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:17 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:16
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.87 | -1.008% | 1.2%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 10.4 | 11.2 | -0.8
Certification Expiration: | 2015-01-21 | 10.4 | 11.3 | -0.9
|  | 10.5 | 11.3 | -0.8
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 745 | 746 | -1 | 631
Certification Expiration: | 2015-01-21 | 745 | 746 | -1 | 41279
|  | 745 | 746 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM
DATE: | 2014-05-09 |  | Model Number: | BAM-1020
TIME: | 14:20 |  | Serial Number: | M4791
0.3
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 14:26 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 14:26
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.35 | 2.141% | -1.92%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22.9 | 24.6 | -1.7
Certification Expiration: | 2015-01-21 | 22.9 | 24.6 | -1.7
|  | 22.9 | 24.6 | -1.7
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 739 | -1 | 638
Certification Expiration: | 2015-01-21 | 738 | 739 | -1 | 41286
|  | 738 | 739 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: NT-SE monitor could not be accessed due to high water surrounding it. |  |  |  |  | Rev. 3 3/2014

## BAM QC NT-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-09 |  | Model Number: | 1020
TIME: | 15:10 |  | Serial Number: | M4793
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 15:16 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 15:15
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.41 | 1.767% | -1.56%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.1 | 22.7 | 0.4
Certification Expiration: | 2015-01-21 | 23.1 | 22.8 | 0.3
|  | 23.1 | 22.7 | 0.4
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 739 | -1 | 638
Certification Expiration: | 2015-01-21 | 738 | 739 | -1 | 41286
|  | 738 | 739 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-09 |  | Model Number: | 1020
TIME: | 14:05 |  | Serial Number: | P21274
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 14:11 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 14:09
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.4 | 16.15 | 1.548% | -3.119%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.5 | 23.5 | 0
Certification Expiration: | 2015-01-21 | 23.5 | 23.5 | 0
|  | 23.6 | 23.5 | 0.1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 737 | 739 | -2 | 638
Certification Expiration: | 2015-01-21 | 737 | 739 | -2 | 41286
|  | 737 | 739 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: Time changed from 14:11 to 14:09.
Rev. 3 3/2014

## BAM QC NT-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-05-09 |  | Model Number: | 1020
TIME: | 15:20 |  | Serial Number: | M4792
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 15:21 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 15:23
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.78 | -0.4768% | 0.6599%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.5 | 23.7 | -0.2
Certification Expiration: | 2015-01-21 | 23.6 | 23.7 | -0.1
|  | 23.5 | 23.6 | -0.1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 739 | 739 | 0 | 638
Certification Expiration: | 2015-01-21 | 739 | 739 | 0 | 41286
|  | 739 | 739 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: Time changed from 14:21 to 14:23.
Rev. 3 3/2014