## BAM QC ST-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-08 |  | Model Number: | BAM-1020
TIME: | 12:30 |  | Serial Number: | P21275
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:31 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:30
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.55 | 0.9063% | -0.7199%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 27.1 | 26.7 | 0.4
Certification Expiration: | 2015-01-21 | 27.1 | 26.7 | 0.4
|  | 27.1 | 26.6 | 0.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 740 | -2 | 637
Certification Expiration: | 2015-01-21 | 738 | 740 | -2 | 41285
|  | 738 | 740 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-08 |  | Model Number: | BAM-1020
TIME: | 10:30 |  | Serial Number: | P22904
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:46 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:45
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.75 | -0.2985% | 0.4799%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.2 | 24.8 | -0.6
Certification Expiration: | 2015-01-21 | 24.2 | 24.6 | -0.4
|  | 24 | 24.7 | -0.7
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 740 | -2 | 637
Certification Expiration: | 2015-01-21 | 738 | 740 | -2 | 41285
|  | 738 | 740 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-08 |  | Model Number: | BAM-1020
TIME: | 10:05 |  | Serial Number: | P22809
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:08 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:07
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.89 | -1.125% | 1.32%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.5 | 24.8 | -0.3
Certification Expiration: | 2015-01-21 | 24.5 | 24.6 | -0.1
|  | 24.5 | 24.7 | -0.2
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 737 | 739 | -2 | 638
Certification Expiration: | 2015-01-21 | 737 | 739 | -2 | 41286
|  | 737 | 739 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-N
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-N |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-08 |  | Model Number: | BAM-1020
TIME: | 13:20 |  | Serial Number: | P22906
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 13:23 |  |  | Certification Expiration: | 2015-01-21 | -100%
Reference Time: | 13:22
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.43 | 1.643% | -1.44%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.7 | 27.3 | -0.6
Certification Expiration: | 2015-01-21 | 26.8 | 27.3 | -0.5
|  | 26.9 | 27.5 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 739 | 740 | -1 | 637
Certification Expiration: | 2015-01-21 | 739 | 740 | -1 | 41285
|  | 739 | 740 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-08 |  | Model Number: | BAM-1020
TIME: | 13:00 |  | Serial Number: | P22808
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 13:02 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 13:01
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.93 | -1.359% | 1.56%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.2 | 27 | -0.8
Certification Expiration: | 2015-01-21 | 26.2 | 27.1 | -0.9
|  | 26.1 | 26.7 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 738 | 740 | -2 | 637
Certification Expiration: | 2015-01-21 | 738 | 740 | -2 | 41285
|  | 738 | 740 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM
DATE: | 2014-07-11 |  | Model Number: | BAM-1020
TIME: | 12:20 |  | Serial Number: | M4791
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:21 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:21
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.65 | 0.3003% | 1.524%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22 | 25.6 | -3.6
Certification Expiration: | 2015-01-21 | 22 | 25.5 | -3.5
|  | 21.9 | 25.2 | -3.3
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: Adjusted temperature from 22.0 to 25.2.
Rev. 3 3/2014

## BAM QC NT-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-11 |  | Model Number: | 1020
TIME: | 13:10 |  | Serial Number: | M4793
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 13:13 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 13:13
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.74 | -0.2389% | 0.4199%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.6 | 25.8 | -1.2
Certification Expiration: | 2015-01-21 | 24.6 | 26 | -1.4
|  | 24.5 | 26 | -1.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 751 | -1 | 626
Certification Expiration: | 2015-01-21 | 750 | 751 | -1 | 41274
|  | 750 | 751 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-11 |  | Model Number: | 1020
TIME: | 11:40 |  | Serial Number: | P21274
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:59 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:59
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.4 | 16.21 | 1.172% | -1.159%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.5 | 26.2 | 0.3
Certification Expiration: | 2015-01-21 | 26.4 | 26.1 | 0.3
|  | 26.4 | 26 | 0.4
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-11 |  | Model Number: | 1020
TIME: | 13:30 |  | Serial Number: | M4792
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 13:33 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 13:33
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.46 | 1.458% | 0.3659%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 26.7 | 27.1 | -0.4
Certification Expiration: | 2015-01-21 | 26.7 | 27 | -0.3
|  | 26.6 | 26.7 | -0.1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 751 | -1 | 626
Certification Expiration: | 2015-01-21 | 750 | 751 | -1 | 41274
|  | 750 | 751 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014