## BAM QC ST-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 10:35 |  | Serial Number: | P21275
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 10:39 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 10:38
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.68 | 0.1199% | 0.05999%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 20.9 | 20.3 | 0.6
Certification Expiration: | 2015-01-21 | 20.9 | 20.1 | 0.8
|  | 20.9 | 20.2 | 0.7
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 750 | -1 | 627
Certification Expiration: | 2015-01-21 | 749 | 750 | -1 | 41275
|  | 749 | 750 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 9:55 |  | Serial Number: | P22904
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:55 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:56
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.84 | -0.8314% | 1.02%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18.4 | 18.6 | -0.2
Certification Expiration: | 2015-01-21 | 18.3 | 18.8 | -0.5
|  | 18.4 | 18.7 | -0.3
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 751 | -1 | 626
Certification Expiration: | 2015-01-21 | 750 | 751 | -1 | 41274
|  | 750 | 751 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 9:40 |  | Serial Number: | P22809
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:44 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:44
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.93 | -1.359% | 1.56%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 19.1 | 18.8 | 0.3
Certification Expiration: | 2015-01-21 | 19.1 | 19.1 | 0
|  | 19.1 | 19.3 | -0.2
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-N
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-N |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 9:25 |  | Serial Number: | P22906
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.4 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:27 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:27
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.54 | 0.9674% | -0.7798%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.9 | 18.7 | -0.8
Certification Expiration: | 2015-01-21 | 17.8 | 18.6 | -0.8
|  | 18 | 18.7 | -0.7
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 750 | 0 | 627
Certification Expiration: | 2015-01-21 | 750 | 750 | 0 | 41275
|  | 750 | 750 | 0 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC ST-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX ST-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 9:10 |  | Serial Number: | P22808
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 9:16 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 9:15
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 17.04 | -1.995% | 2.22%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18.2 | 19.1 | -0.9
Certification Expiration: | 2015-01-21 | 18.2 | 19 | -0.8
|  | 18.2 | 19.1 | -0.9
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 750 | -1 | 627
Certification Expiration: | 2015-01-21 | 749 | 750 | -1 | 41275
|  | 749 | 750 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-SE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM
DATE: | 2014-07-24 |  | Model Number: | BAM-1020
TIME: | 11:30 |  | Serial Number: | M4791
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:44 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:44
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.74 | -0.2389% | 2.073%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.1 | 20.2 | -3.1
Certification Expiration: | 2015-01-21 | 17.2 | 20.3 | -3.1
|  | 17.1 | 20.2 | -3.1
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 751 | -2 | 626
Certification Expiration: | 2015-01-21 | 749 | 751 | -2 | 41274
|  | 749 | 751 | -2 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Note: Adjusted temperature from 17.2 to 20.0.
Rev. 3 3/2014

## BAM QC NT-SW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-SW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | 1020
TIME: | 12:00 |  | Serial Number: | M4793
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:12 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:12
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.68 | 0.1199% | 0.05999%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 19 | 20.2 | -1.2
Certification Expiration: | 2015-01-21 | 19.2 | 20.7 | -1.5
|  | 19.1 | 20.6 | -1.5
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 751 | -1 | 626
Certification Expiration: | 2015-01-21 | 750 | 751 | -1 | 41274
|  | 750 | 751 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NE
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | 1020
TIME: | 11:00 |  | Serial Number: | P21274
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.3 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 11:02 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 11:02
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.4 | 16.33 | 0.4287% | -0.4268%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23 | 23.7 | -0.7
Certification Expiration: | 2015-01-21 | 23.2 | 23.8 | -0.6
|  | 23.1 | 23.7 | -0.6
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 749 | 750 | -1 | 627
Certification Expiration: | 2015-01-21 | 749 | 750 | -1 | 41275
|  | 749 | 750 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014

## BAM QC NT-NW
BAM-1020 PARTICULATE MATTER (PM) MONITOR
Calibration Data
SITE NAME: | KCBX NT-NW |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | BAM-1020
DATE: | 2014-07-24 |  | Model Number: | 1020
TIME: | 12:20 |  | Serial Number: | M4792
FLOW RATE AND LEAK CHECK Calibration Data
Leak Check1: | 0.5 | LPM |  | Flow Rate Calibration Device: | BGI
1. Acceptance Criteria: < 0.5 LPM |  |  |  | Model Number: | deltaCal
Clock/Timer Verification2: |  |  |  | Serial Number: | 1377
BAM-1020 : | 12:20 |  |  | Certification Expiration: | 2015-01-21
Reference Time: | 12:20
2. Acceptance Criteria: within 60 sec. of Reference Time
|  | Sampler Indicated | Calibration | Percent Difference3 | Percent Difference4
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 16.7 | 16.49 | 1.273% | 0.5488%
The BAM-1020 operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 3. Acceptance Criteria: ± 4.0%
4. Acceptance Criteria: ± 5.0%
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference5
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 22.6 | 22 | 0.6
Certification Expiration: | 2015-01-21 | 22.6 | 22 | 0.6
|  | 22.4 | 21.6 | 0.8
5. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference6
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 751 | -1 | 626
Certification Expiration: | 2015-01-21 | 750 | 751 | -1 | 41274
|  | 750 | 751 | -1 | --
6. Acceptance Criteria: ± 10.0 mm Hg
Calibration Tech: | Greg Mazik
Rev. 3 3/2014