## PM Sampler Checklist North
KCBX TERMINALS CO NORTH (NT-NE) | SITE VISIT DATE
| 2014-06-24 | 2014-06-27
PartisolPlus 2025i Metals/Mass Sampler
Arrival time: | 9:00 | 8:10
Any status flags on main screen? | N | N
Is sampler date/time correct within 60 sec? | Y | Y
Have data from last run been collected? | N | N
Has sampled filter magazine been collected? | N | Y
Has COC form for retrieved samples been filled out? | N | N
Have samples been packed for shipping? | N | N
Haves samples been shipped to lab? | N | N
Have sample data been sent to Austin staff? | N | N
Is cassette magazine in good condition? | Y | Y
Are cassettes available in feed magazine for next sample? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
PartisolPlus 2025i Carbon Sampler
Arrival time: | 9:00 | 8:10
Any status flags on main screen? | N | N
Is sampler date/time correct within 60 sec? | Y | Y
Have data from last run been collected? | N | N
Has sampled filter magazine been collected? | N | Y
Has COC form for retrieved samples been filled out? | N | N
Have samples been packed for shipping? | N | N
Haves samples been shipped to lab? | N | N
Have sample data been sent to Austin staff? | N | N
Is cassette magazine in good condition? | Y | Y
Are cassettes available in feed magazine for next sample? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
Comments section
6/27, removed cassettes that ran on 6/16, 6/19, and 6/22.
MAINTENANCE LIST FOR 2025i
| Each Visit: | Inspect filter cassettes and magazines; discard damaged cassettes, and lubricate cassette seals as needed
| Weekly: | Perform external leak check | Rev 1 (1/2014)
| Monthly: | Perform internal leak check
Clean PM10 inlet
| Biannually: | Change the large in-line filter
Clean the rain hoods and air screens
| Annually: | Check and rebuild pump if needed

## PM Sampler Checklist South
KCBX TERMINALS CO SOUTH (ST-NE) | SITE VISIT DATE
| 2014-06-24 | 2014-06-27
PartisolPlus 2025i Metals/Mass Sampler
Arrival time: | 10:30 | 11:30
Any status flags on main screen? | N | N
Is sampler date/time correct within 60 sec? | N | Y
Have data from last run been collected? | N | N
Has sampled filter magazine been collected? | N | Y
Has COC form for retrieved samples been filled out? | N | N
Have samples been packed for shipping? | N | N
Haves samples been shipped to lab? | N | N
Have sample data been sent to Austin staff? | N | N
Is cassette magazine in good condition? | Y | Y
Are cassettes available in feed magazine for next sample? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
PartisolPlus 2025i Carbon Sampler
Arrival time: | 10:30 | 11:30
Any status flags on main screen? | N | N
Is sampler date/time correct within 60 sec? | Y | Y
Have data from last run been collected? | N | N
Has sampled filter magazine been collected? | N | Y
Has COC form for retrieved samples been filled out? | N | N
Have samples been packed for shipping? | N | N
Haves samples been shipped to lab? | N | N
Have sample data been sent to Austin staff? | N | N
Is cassette magazine in good condition? | Y | Y
Are cassettes available in feed magazine for next sample? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
Comments section
6/24, changed time on ST-metals Partisol from 11:18
to 11:20.
6/27, removed cassettes that ran on 6/16, 6/19, and 6/22.
MAINTENANCE LIST FOR 2025i
| Each Visit: | Inspect filter cassettes and magazines; discard damaged cassettes, and lubricate cassette seals as needed
| Weekly: | Perform external leak check | Rev 1 (1/2014)
| Monthly: | Perform internal leak check
Clean PM10 inlet
| Biannually: | Change the large in-line filter
Clean the rain hoods and air screens
| Annually: | Check and rebuild pump if needed

## MetBAM Checklist North
KCBX TERMINALS CO NORTH | SITE VISIT DATE
| 2014-06-24 | 2014-06-27
Meteorological Station
Arrival time: | 9:00 | 8:10
Is the temperature aspirator fan running? | Y | Y
Are all signal cables secured and intact? | Y | Y
Are wind vane/ anemometer cups intact and operational? | Y | Y
Is datalogger date/time correct within 60 sec? | NA | NA
Are conditions shown on logger consistent with ambient observations? | NA | NA
BAM-1020 PM10 Monitors
Arrival time: | 9:00 | 8:10
Any status error flags showing? | N | N
Are all monitor dates/times correct within 60 sec? | Y | Y
Are all filter transport systems operational? | Y | Y
Are all tape reel cover knobs secure? | Y | Y
Are hourly data being collected? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
Comments section
Site ID List: | MAINTENANCE LIST FOR METEOROLOGICAL STATION
NT-NW = Northwest BAM site (Meteorological monitors) | Each Visit: | Inspect sensors and cables for signs of deterioration or damage and replace as needed
NT-NE = Northeast BAM site (Partisol samplers) | Monthly: | Inspect temperature aspirator and remove dust accumulation | Rev 3 (5/2014)
NT-SE = Southeast BAM site |  | Inspect tower for plumb, and verify that all fasteners are intact and tight
NT-SW = Southwest BAM site | Biannually: | Perform calibration check on sensors
Check wind sensor bearings
| Annually: | Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
| Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
| Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
| Quarterly: | Verify BAM settings file
Flow system Calibration
Disassemble and clean inlet
| Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
| Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring

## MetBAM Checklist South
KCBX TERMINALS CO SOUTH | SITE VISIT DATE
| 2014-06-24 | 2014-06-27
Meteorological Station
Arrival time: | 10:30 | 11:30
Is the temperature aspirator fan running? | Y | Y
Are all signal cables secured and intact? | Y | N
Are wind vane/ anemometer cups intact and operational? | Y | Y
Is datalogger date/time correct within 60 sec? | NA | NA
Are conditions shown on logger consistent with ambient observations? | NA | NA
BAM-1020 PM10 Monitors
Arrival time: | 10:30 | 11:30
Any status error flags showing? | N | N
Are all monitor dates/times correct within 60 sec? | Y | Y
Are all filter transport systems operational? | Y | Y
Are all tape reel cover knobs secure? | Y | Y
Are hourly data being collected? | Y | Y
Is sampler inlet and downtube clean? | Y | Y
Is water collection bottle clean? | Y | Y
Are O-rings in good condition? | Y | Y
Comments section
6/27, lightning rod cable loose. Secured it to the signal cable conduit.
Site ID List: | MAINTENANCE LIST FOR METEOROLOGICAL STATION
ST-NW = Northwest BAM site | Each Visit: | Inspect sensors and cables for signs of deterioration or damage and replace as needed
ST-N = North BAM site | Monthly: | Inspect temperature aspirator and remove dust accumulation | Rev 3 (5/2014)
ST-NE = Northeast BAM site (Partisol samplers) |  | Inspect tower for plumb, and verify that all fasteners are intact and tight
ST-M = Meteorological Monitoring Site | Biannually: | Perform calibration check on sensors
ST-CE = Center East BAM site |  | Check wind sensor bearings
ST-SW = Southwest BAM site | Annually: | Replace wind sensor bearings
MAINTENANCE LIST FOR BAM-1020
| Each Visit: | Inspect filter tape and replace when needed (~2 months/roll; run SELF-TEST function in the TAPE menu).
Inspect filter tape for pinholes and irregular PM deposits
| Monthly: | Perform flow rate and leak check
Clean BAM nozzle and vane
Clean inlet hardware, capstan shaft and pinch roller
| Quarterly: | Verify BAM settings file
Flow system calibration
Disassemble and clean inlet
| Biannually: | Check pump muffler and tubing and replace as needed
Test filter RH and Temperature sensors
Test smart heater function
| Annually: | Perform zero background test
Clean internal debris filter
Remove and check membrane span foil
Beta detector count rate and dark count test
Rebuild vacuum pump
Replace nozzle o-ring