## ST-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Metals
DATE: | 2014-05-09 |  | Model Number: | 2025i
TIME: | 12:50 |  | Serial Number: | 20547
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.4 | 23.6 | 0.8
Certification Expiration: | 2015-01-21 | 24.6 | 23.6 | 1
|  | 24.5 | 23.5 | 1
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 737 | 739 | -2 | 638
Certification Expiration: | 2015-01-21 | 737 | 739 | -2 | 41286
|  | 737 | 739 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.3 | 23.7 | -0.4
Certification Expiration: | 2015-01-21 | 23.3 | 23.7 | -0.4
|  | 23.3 | 23.8 | -0.5
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 5 | mmHg | Partisol : | 12:56
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 12:56
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 16.95 | -1.652% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## ST-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Carbon
DATE: | 2014-05-09 |  | Model Number: | 2025i
TIME: | 12:30 |  | Serial Number: | 20544
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.5 | 24.2 | 0.3
Certification Expiration: | 2015-01-21 | 24.5 | 24.1 | 0.4
|  | 24.6 | 24.1 | 0.5
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 739 | 738 | 1 | 639
Certification Expiration: | 2015-01-21 | 739 | 738 | 1 | 41287
|  | 739 | 738 | 1 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.2 | 23.7 | -0.5
Certification Expiration: | 2015-01-21 | 23.2 | 23.7 | -0.5
|  | 23.3 | 23.8 | -0.5
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 13 | mmHg | Partisol : | 12:46
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 12:46
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.67 | 16.31 | 2.207% | 0%
Certification Expiration: | 2015-01-21 | 16.67 | 16.31 | 2.207% | 0%
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Metals
DATE: | 2014-05-09 |  | Model Number: | 2025i
TIME: | 13:35 |  | Serial Number: | 20546
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.7 | 23.8 | -0.1
Certification Expiration: | 2015-01-21 | 23.8 | 23.8 | 0
|  | 23.8 | 23.8 | 0
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 737 | 739 | -2 | 638
Certification Expiration: | 2015-01-21 | 737 | 739 | -2 | 41286
|  | 737 | 739 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 25 | 25.3 | -0.3
Certification Expiration: | 2015-01-21 | 25 | 25.2 | -0.2
|  | 25 | 25.2 | -0.2
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 10 | mmHg | Partisol : | 13:44
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 13:44
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.68 | 17.07 | -2.285% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Carbon
DATE: | 2014-05-09 |  | Model Number: | 2025i
TIME: | 13:50 |  | Serial Number: | 20545
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 23.3 | 24 | -0.7
Certification Expiration: | 2015-01-21 | 23.3 | 23.8 | -0.5
|  | 23.3 | 23.7 | -0.4
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 737 | 739 | -2 | 638
Certification Expiration: | 2015-01-21 | 737 | 739 | -2 | 41286
|  | 737 | 739 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 24.6 | 25.3 | -0.7
Certification Expiration: | 2015-01-21 | 24.6 | 25.4 | -0.8
|  | 24.6 | 25.4 | -0.8
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
|  |  | Clock/Timer Verification5: | 11:37
External Leak Check4 | 10 | mmHg | Partisol : | 13:51
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 13:51
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 | 16.68 | 16.51 | 1.03% | 0%
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM