## ST-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Metals
DATE: | 2014-05-22 |  | Model Number: | 2025i
TIME: | 11:50 |  | Serial Number: | 20547
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 16.3 | 15.6 | 0.7
Certification Expiration: | 2015-01-21 | 16.3 | 15.6 | 0.7
|  | 16.3 | 15.6 | 0.7
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 748 | 750 | -2 | 627
Certification Expiration: | 2015-01-21 | 748 | 750 | -2 | 41275
|  | 748 | 750 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 18.2 | 19.2 | -1
Certification Expiration: | 2015-01-21 | 18.3 | 19.2 | -0.9
|  | 18.2 | 19.2 | -1
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 5 | mmHg | Partisol : | 11:56
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 11:56
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377 |  | 16.95
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.09 | 15.53 | -2.833% | 0%
|  | 18.39 | 18.9 | -2.698% | 0%
|  | 16.66 | 16.88 | -1.303% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## ST-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX ST-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | ST-NE-Carbon
DATE: | 2014-05-22 |  | Model Number: | 2025i
TIME: | 11:25 |  | Serial Number: | 20544
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 16.4 | 15.6 | 0.8
Certification Expiration: | 2015-01-21 | 16.7 | 15.6 | 1.1
|  | 16.6 | 15.6 | 1
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 750 | 749 | 1 | 628
Certification Expiration: | 2015-01-21 | 750 | 749 | 1 | 41276
|  | 750 | 749 | 1 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.9 | 16.7 | 1.2
Certification Expiration: | 2015-01-21 | 17.9 | 16.7 | 1.2
|  | 17.9 | 16.7 | 1.2
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 23 | mmHg | Partisol : | 11:30
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 11:30
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.1 | 14.84 | 1.752% | 0%
|  | 18.41 | 18.36 | 0.2723% | 0%
|  | 16.67 | 16.57 | 0.6035% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Metals
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Metals
DATE: | 2014-05-22 |  | Model Number: | 2025i
TIME: | 8:10 |  | Serial Number: | 20546
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 15.2 | 15.1 | 0.1
Certification Expiration: | 2015-01-21 | 15.3 | 15.2 | 0.1
|  | 15.3 | 14.9 | 0.4
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 747 | 749 | -2 | 628
Certification Expiration: | 2015-01-21 | 747 | 749 | -2 | 41276
|  | 747 | 749 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.7 | 17 | 0.7
Certification Expiration: | 2015-01-21 | 17.8 | 17 | 0.8
|  | 17.8 | 17 | 0.8
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
Clock/Timer Verification5:
External Leak Check4 | 21 | mmHg | Partisol : | 8:16
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 8:16
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.1 | 15 | 0.6667% | 0%
|  | 18.4 | 18.39 | 0.05438% | 0%
|  | 16.7 | 16.96 | -1.533% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM

## NT-NE-Carbon
PARTISOL PARTICULATE MATTER (PM) SAMPLER
Calibration Data
SITE NAME: | KCBX NT-NE |  | CLIENT: | Koch
AQS Site Code: | N/A |  | SAMPLER ID: | NT-NE-Carbon
DATE: | 2014-05-22 |  | Model Number: | 2025i
TIME: | 8:30 |  | Serial Number: | 20545
TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Temperature Difference1
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 15 | 15.1 | -0.1
Certification Expiration: | 2015-01-21 | 15 | 15.1 | -0.1
|  | 15.1 | 15.1 | 0
1. Acceptance Criteria: ± 2.0 °C
PRESSURE Calibration Data
Calibration Device: | BGI | Sampler Sensor | Calibration Sensor | Pressure Difference2
Model Number: | deltaCal | Pressure (Pa) (mm Hg) | Pressure (Pa) (mm Hg) | (Sampler - Calibration) (mm Hg)
Serial Number: | 1377 | 748 | 750 | -2 | 627
Certification Expiration: | 2015-01-21 | 748 | 750 | -2 | 41275
|  | 748 | 750 | -2 | --
2. Acceptance Criteria: ± 10.0 mm Hg
COMPARTMENT TEMPERATURE Calibration Data
Calibration Device: | BGI | Sampler Compartment | Calibration Sensor | Temperature Difference3
Model Number: | deltaCal | Temperature (Ta) (°C) | Temperature (Ta) (°C) | (Sampler - Calibration) (°C)
Serial Number: | 1377 | 17.4 | 16.3 | 1.1
Certification Expiration: | 2015-01-21 | 17.4 | 16.3 | 1.1
|  | 17.4 | 16.4 | 1
3. Acceptance Criteria: ± 2.0 °C
LEAK CHECK, TIME, AND FLOWRATE Calibration Data
|  |  | Clock/Timer Verification5: | 11:37
External Leak Check4 | 13 | mmHg | Partisol : | 8:40
| 4. Acceptance Criteria < 25 mm Hg |  | Reference Time: | 8:42
5. Acceptance Criteria: within 60 sec. of Reference Time
SINGLE POINT CHECK
Calibration Device: | BGI | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
Model Number: | deltaCal | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
Serial Number: | 1377
Certification Expiration: | 2015-01-21 | THREE POINT CHECK
|  | Sampler Indicated | Calibration | Percent Difference6 | Percent Difference7
|  | Flow Rate (LPM) | Flow Rate (Qa) (LPM) | (Sampler vs. Calibration) | (Calibration vs. Design)
|  | 15.09 | 14.95 | 0.9365% | 0%
|  | 18.4 | 18.36 | 0.2179% | 0%
|  | 16.68 | 16.56 | 0.7246% | 0%
A three point flow check is required quarterly and a single point operating flow rate check is required monthly |  |  |  | 6. Acceptance Criteria: ± 4.0%
The 2025i Partisol operates using actual flow rates, and converts to standard conditions in the onboard data file |  |  |  | 7. Acceptance Criteria: ± 5.0%
Calibration Tech: | G. Mazik
Note: Changed time from 08:40 to 08:42.
Rev. 3 3/2014
Note on D41: Andy Hodgson: Setpoint at 15LPM
Note on D45: Andy Hodgson: Setpoint at 15LPM
Note on D46: Andy Hodgson: Setpoint at 15LPM
Note on D47: Andy Hodgson: Setpoint at 15LPM