## Readme
INSTRUCTIONS FOR MATERIAL WORKBOOK
Purpose
This worksheet supports linear, quadratic, cubic, and quartic regression models:
| Linear: | y = b0 + b1x+ e
| Quadratic: | y = b0 + b1x + b2x2 + e
| Cubic: | y = b0 + b1x + b2x2 + b3x3 +e
| Quartic: | y = b0 + b1x + b2x2 + b3x3 +b4x4 + e
However, it will not perform the Zero Air Test for the cubic or quartic models.
The inputs are:
|  | x | the concentration of the certified calibration standard
|  | y | the measured instrument response
|  | e | a random effect due to variation of influence quantities
It is assumed that the concentrations of the calibration standards (x) have negligible uncertainties.
The output quantities are estimated parameters (bs) and related uncertainties.
The workbook then helps the user perform the following functions:
- | determine which model (linear, quadratic, etc.) is best.
- | determine whether zero and span responses are acceptable.
- | based upon maximum concentration allowed give the maximum acceptable maximum response
analyzed on the same day as the initial calibration or a subsequent day
Warning
EPA discourages the use of cubic and quartic models and believes that better fits of the data
can be obtained by performing multipoint calibrations over more limited concentration ranges
and by using straight-line or quadratic models. Inclusion of cubic and quartic models in the
spreadsheet is for experimental use or for situations in which there is a theoretical basis for
the use of such higher-order models. Analysts should be aware that apparent higher-order
calibration curves may be caused by artifacts such as inaccurate reference standards or
leaks in a gas dilution system. They should not use higher-order regression models to fit
multipoint calibration data that have inadequate precision and that should be fitted to lower-
order regression models. If analysts suspect that precision is inadequate, they should make
replicate measurements at each different concentration. Additionally, a multipoint calibration
should not change regression model orders from one month to the next.
Organization
The workbook consists of several worksheets, which are displayed as tabs at the bottom of
the screen. The functions of these worksheets are described below:
ReadMe (this sheet) - describes the workbook, explaining how to use the worksheets
Measurement Data - allows for user input of calibration and other analytical data
and includes statistical calculations for polynomial regression. At the end it will give
the maximum measurement that can be observed in order to consider the gas a zero
air material.
Curves 1 - displays the calibration data, the best-fit line, and expanded uncertainties.
Residuals 1 - displays the difference between the observed responses and those
estimated by the best-fit calibration line.
Curves 2 - displays the calibration data, the best-fit quadratic curve, and
expanded uncertainties.
Residuals 2 - displays the difference between the observed responses and those
estimated by the best-fit quadratic regression line.
Curves 3 - displays the calibration data, the best-fit cubic curve, and
expanded uncertainties.
Residuals 3 - displays the difference between the observed responses and those
estimated by the best fit cubic regression line.
Curves 4 - displays the calibration data, the best-fit quartic curve, and
expanded uncertainties.
Residuals 4 - displays the difference between the observed responses and those
estimated by the best-fit quartic regression line.
Chart Data - includes the data used to create the Curves and Residuals charts
Conventions
The Zero Test worksheet includes instructions that guide the user through the
steps in its use. The worksheet is also color coded to simplify use. Shaded cells that are
bordered in blue lines are for input of data. These cells are unprotected, but all other cells of
the Zero Test worksheet are protected. The only other unprotected cell in the
workbook is cell F4 of the Chart Data worksheet. That cell controls the approximate level of
confidence associated with the expanded uncertainties displayed in the "Curves" charts.
Derived values and statements are colored red. These cells contain formulas and are
protected to prevent alteration.
Spreadsheet background colors indicate the order of the polynomial supported by the
calculations in the area.
Light green is used for the linear model.
Yellow is used for the quadratic model.
Gray color is used for the cubic model.
Light blue is used for the quartic model.
Use
The Measurement Data worksheet guides the user through six steps.
STEP 1 |  |  | Enter Calibration Data
In this step, up to 50 calibration points may be entered. Each calibration point has two parts, the
certified concentration of the calibration gas standard and the instrument response when testing
the standard. These values are entered in two columns. The spreadsheet performs
computations in cloumns I through P (linear), Q through X (quadratic), Y through AZ (cubic), and BA
and above (quartic).
STEP 2 |  |  | Review the Parameter Estimates
In this step, the user reviews the estimates of the intercepts (b0), slopes (b1) and other
parameters (b2, b3, and b4) for the four models, examines their expanded uncertainties and the
overall uncertainty of the fit (u2). The result of an F test indicates which of the models is best. The
linear model is recommended unless the quadratic or higher order model significantly reduces
the residual error.
STEP 3 |  | Review the Charts
In this step, the user reviews the charts named Curves 1, Residuals 1, Curves 2, Residuals 2, etc.
These charts help the user understand why one model performs better than the other.
STEP 4 |  | Test Zero Air Material on Same Day as Initial Calibration
In this step, the user enters the maximum concentration for the gas calibrated and being measured
so that it will be considered a zero air material. Based upon the calibration of the instrument
they will get what the maximum measurement can be in order to consider this sufficiently
less than the maximum concentration at a 95% level.
STEP 5 |  | Test Zero Air Material on Different Day of Initial Calibration
In this step, the user enters the responses to a candidate that is tested on a different day
from the calibration of STEP 1. The worksheet first assesses the zero and span responses. If
the zero and span responses are acceptable, the user proceeds to enter the maximum acceptable
concentration. The results include those for zero and span reference standards. The
table below specifies the number of different reference span standards concentrations needed for
each model, in addition to the zero standard. Failure to meet these requirements will result in serious
calculation errors and invalid results.
| Model | Minimum number of different nonzero reference standards required
| Linear | 1
| Quadratic | 2
| Cubic | NA
| Quartic | NA
The spreadsheet determines whether the regression curve has changed since the initial
calibration. The data are corrected for any change and the estimated concentration of the
candidate standard is provided together with its expanded uncertainty associated with an
approximate level of confidence of 95%.
After confirming and adjusting for any changes, the user can enter the maximum allowable concentration
for the gas, and will get the maximum allowable response in order to confirm it is a zero material at a
95% level.
Reference: Appendix A, EPA Traceability Protocol Assay and Certification of Gaseous
Calibration Standards, USEPA, 1997.

## Zero Test
This spreadsheet is used for the Zero Material test.
| STEP 1 |  | Enter Calibration Data | [8 empty] | Calculations - Linear Model |  |  |  |  | Calculations - Quadratic Model | [6 empty] | Calculations - Cubic Model | [5 empty] | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model |  | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model
Enter the calibration data in the shaded spaces below. The first column (i)
simply counts the calibration points that you enter. The second column (X) is
for the certified concentrations of the calibration gas standards. The third
column (Y) is for the instrument responses corresponding to the calibration
standards. The number of points cannot exceed 50.
| i | Xi | Yi |  | Color Code | [6 empty] | X'X |  |  |  |  | X2 |  | X'X |  |  |  |  | X3 | X'X | [7 empty] | X4 |  | X'X
| 1 | 0 | 0.19 | [8 empty] | 18 | 45 |  |  |  | 0 |  | 18 | 45 | 165 |  |  | 0 | 18 | 45 | 165 | 675 |  |  |  |  | 0 |  | 18 | 45 | 165 | 675 | 2937
| 2 | 0 | 0.2 |  | red = derived value (protected) | [6 empty] | 45 | 165 |  |  |  | 0 |  | 45 | 165 | 675 |  |  | 0 | 45 | 165 | 675 | 2937 |  |  |  |  | 0 |  | 45 | 165 | 675 | 2937 | 13275
| 3 | 0 | 0.21 | [13 empty] | 0 |  | 165 | 675 | 2937 |  |  | 0 | 165 | 675 | 2937 | 13275 |  |  |  |  | 0 |  | 165 | 675 | 2937 | 13275 | 61545
| 4 | 1 | 1.14 |  | blue = entered value (unprotected) | [6 empty] | Det(X'X) |  |  |  |  | 1 | [6 empty] | 1 | 675 | 2937 | 13275 | 61545 |  |  |  |  | 1 |  | 675 | 2937 | 13275 | 61545 | 290475
| 5 | 1 | 1.15 | [8 empty] | 945 |  |  |  |  | 1 |  | Det(X'X) |  |  |  |  | 1 | Det(X'X) | [7 empty] | 1 |  | 2937 | 13275 | 61545 | 290475 | 1388937
| 6 | 1 | 1.16 |  | black = fixed text (protected) | [11 empty] | 1 |  | 105840 |  |  |  |  | 1 | 20575296 | [7 empty] | 1 | 1 | 2 | 3 | 4 | 5
| 7 | 2 | 1.99 | [8 empty] | X'Xinverse |  |  |  |  | 4 | [6 empty] | 8 | [8 empty] | 1: 16; 1: 2; 2: X'Xinverse
| 8 | 2 | 2 | [8 empty] | 0.1746 | -0.047619 |  |  |  | 4 |  | X'Xinverse |  |  |  |  | 8 | X'Xinverse | [7 empty] | 16 | 3 | 0.33201 | -0.60626 | 0.3588 | -0.084877 | 0.0069444
| 9 | 2 | 2.01 | [8 empty] | -0.047619 | 0.019048 |  |  |  | 4 |  | 0.27381 | -0.19643 | 0.029762 |  |  | 8 | 0.32011 | -0.40785 | 0.1455 | -0.015432 |  |  |  |  | 16 | 4 | -0.60626 | 4.5146 | -4.1281 | 1.2279 | -0.11574
| 10 | 3 | 2.74 | [13 empty] | 9 |  | -0.19643 | 0.24226 | -0.044643 |  |  | 27 | -0.40785 | 1.2077 | -0.57319 | 0.070473 |  |  |  |  | 81 | 5 | 0.3588 | -4.1281 | 4.1198 | -1.2828 | 0.12442
| 11 | 3 | 2.75 | [8 empty] | X'Y |  |  |  |  | 9 |  | 0.029762 | -0.044643 | 0.0089286 |  |  | 27 | 0.1455 | -0.57319 | 0.29828 | -0.03858 |  |  |  |  | 81 |  | -0.084877 | 1.2279 | -1.2828 | 0.41024 | -0.040509
| 12 | 3 | 2.76 | [8 empty] | 40.35 |  |  |  |  | 9 | [6 empty] | 27 | -0.015432 | 0.070473 | -0.03858 | 0.005144 |  |  |  |  | 81 |  | 0.0069444 | -0.11574 | 0.12442 | -0.040509 | 0.0040509
| 13 | 4 | 3.39 | [8 empty] | 140.25 |  |  |  |  | 16 |  | X'Y |  |  |  |  | 64 | X'Y | [7 empty] | 1: 256
| 14 | 4 | 3.4 | [13 empty] | 16 |  | 40.35 |  |  |  |  | 64 | 40.35 | [7 empty] | 1: 256; 2: X'Y; 4: b
| 15 | 4 | 3.41 | [8 empty] | b |  | b' |  |  | 16 |  | 140.25 |  |  |  |  | 64 | 140.25 | [7 empty] | 1: 256; 2: 40.35; 4: 0.2
| 16 | 5 | 3.94 | [8 empty] | 0.36667 |  | 0.36667 | 0.75 |  | 25 |  | 561.15 |  |  |  |  | 125 | 561.15 | [7 empty] | 1: 625; 2: 140.25; 4: 1
| 17 | 5 | 3.95 | [8 empty] | 0.75 |  |  |  |  | 25 | [6 empty] | 125 | 2408.25 | [7 empty] | 1: 625; 2: 561.15; 4: -0.05
| 18 | 5 | 3.96 | [13 empty] | 25 |  | b |  | b' |  |  | 125 | b | [7 empty] | 1: 625; 2: 2408.25; 4: 1.7053e-13
[12 empty] | Y'Y | [6 empty] | 0.2 |  | 0.2 | 1 | -0.05 |  | 0.2 | [7 empty] | 2: 10785.15; 4: -2.1316e-14
[12 empty] | 120.2637 | [6 empty] | 1 | [5 empty] | 1
[19 empty] | -0.05 | [5 empty] | -0.05
[12 empty] | df | [12 empty] | 3.5527e-15
[12 empty] | 16 | [6 empty] | t(0.95) | [5 empty] | t(0.95) | [7 empty] | 2: t(0.95)
[19 empty] | 2.1314 | [5 empty] | 2.1448 | [7 empty] | 2: 2.1604
t(0.95,df)
[12 empty] | 2.1199 | [6 empty] | b'X'Y = SSmodel | [5 empty] | b'X'Y = SSmodel | [7 empty] | 2: b'X'Y = SSmodel
[19 empty] | 120.2625 | [5 empty] | 120.2625 | [7 empty] | 2: 120.2625
b'X'Y = SSmodel
[12 empty] | 119.9825 | [6 empty] | SSregression | [5 empty] | SSregression | [7 empty] | 2: SSregression
[19 empty] | 29.8113 | [5 empty] | 29.8113 | [7 empty] | 2: 29.8113
[12 empty] | SSregression | [14 empty] | b' | bar: 4
[12 empty] | 29.5313 | [6 empty] | SSresidual | [5 empty] | SSresidual |  | 0.2 | 1 | -0.05 | 3.5527e-15 |  |  | 2: SSresidual; 4: 0.2; 5: 1 | -0.05 | 1.7053e-13 | -2.1316e-14
[19 empty] | 0.0012 | [5 empty] | 0.0012 | [7 empty] | 2: 0.0012
SSresidual
[12 empty] | 0.2812 | [6 empty] | s2 | [5 empty] | s2 | bar: 2
[19 empty] | 8e-05 | [5 empty] | 8.5714e-05 | [7 empty] | 2: 9.2308e-05
s2
[12 empty] | 0.017575 | [6 empty] | Fcrit(0.05) | [5 empty] | Fcrit(0.05) | [7 empty] | 2: Fcrit(0.05)
[19 empty] | 2.3849 | [5 empty] | 2.463 | [7 empty] | 2: 2.5536
Var-Covar
[12 empty] | 0.0030687 | -0.0008369 | [5 empty] | F | [5 empty] | F | bar: 2
[12 empty] | -0.0008369 | 0.00033476 | [5 empty] | 219.6875 | [5 empty] | 0.93333 | [7 empty] | 2: 0.92857
[12 empty] | n | [6 empty] | Preferred Model | [5 empty] | Preferred Model | [7 empty] | 2: Preferred Model
[12 empty] | 18 | [6 empty] | Q | [5 empty] | Quadratic | [7 empty] | 2: Cubic
[19 empty] | Var-Covar | [5 empty] | Var-Covar | [7 empty] | 2: Var-Covar
[19 empty] | 2.1905e-05 | -1.5714e-05 | 2.381e-06 |  |  |  | 2.5608e-05 | -3.2628e-05 | 1.164e-05 | -1.2346e-06 | [6 empty] | 3.0647e-05 | -5.5963e-05 | 3.312e-05 | -7.8348e-06 | 6.4103e-07
[19 empty] | -1.5714e-05 | 1.9381e-05 | -3.5714e-06 |  |  |  | -3.2628e-05 | 9.662e-05 | -4.5855e-05 | 5.6379e-06 | [6 empty] | -5.5963e-05 | 0.00041673 | -0.00038105 | 0.00011334 | -1.0684e-05
[19 empty] | 2.381e-06 | -3.5714e-06 | 7.1429e-07 |  |  |  | 1.164e-05 | -4.5855e-05 | 2.3862e-05 | -3.0864e-06 | [6 empty] | 3.312e-05 | -0.00038105 | 0.00038029 | -0.00011841 | 1.1485e-05
[25 empty] | -1.2346e-06 | 5.6379e-06 | -3.0864e-06 | 4.1152e-07 | [6 empty] | -7.8348e-06 | 0.00011334 | -0.00011841 | 3.7868e-05 | -3.7393e-06
| STEP 2 |  | Review the Parameter Estimates | [31 empty] | 6.4103e-07 | -1.0684e-05 | 1.1485e-05 | -3.7393e-06 | 3.7393e-07
Review the estimates of the parameters (b0, b1, …)
for the linear and quadratic models, their expanded
uncertainties and the overall uncertainty of the fit (u2).
The approximate level of confidence
| Linear Model |  | Expanded Uncertainty |  |  | associated with the expanded
| parameter | estimate | lower | upper |  | uncertainties is 95%.
| b0 = | 0.36667 | 0.24923 | 0.4841
| b1 = | 0.75 | 0.71121 | 0.78879 |  | The corresponding intervals take the form
| s2 = | 0.017575 | 0.0097485 | 0.040708 |  | estimate +/- t*u, where:
| s = | 0.13257 | 0.098735 | 0.20176 |  | estimate is the quantity of interest,
| df = | 16 |  |  |  | u is its uncertainty, and
| t = | 2.1199 |  |  |  | t is a coverage factor.
| Quadratic Model |  | Expanded Uncertainty |  |  | For each model,
| parameter | estimate | lower | upper |  | the coverage factor, t, is
| b0 = | 0.2 | 0.19995 | 0.20005 |  | determined from the t-distribution
| b1 = | 1 | 0.99996 | 1 |  | with appropriate degrees of
| b2 = | -0.05 | -0.050002 | -0.049998 |  | freedom (df).
| s2 = | 8e-05 | 4.343e-05 | 0.00019163
| s = | 0.0089443 | 0.0065902 | 0.013843 |  | Note that the uncertainties are not
| df = | 15 |  |  |  | displayed, but can easily be
| t = | 2.1314 |  |  |  | derived from the estimate, expanded
uncertainty, and the coverage factor.
Comparing the two models:
| F ratio = | 219.6875
| F critical = | 2.3849 | (5% significance level)
The quadratic model produces a significantly smaller overall uncertainty of fit.
The quadratic model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then review the following estimates
for those models. Otherwise, go to Step 3.
| Cubic Model |  | Expanded Uncertainty |  |  | See above notes on calculation
| parameter | estimate | lower | upper |  | of expanded uncertainties.
| b0 = | 0.2 | 0.18915 | 0.21085
| b1 = | 1 | 0.97892 | 1.0211
| b2 = | -0.05 | -0.060477 | -0.039523
| b3 = | 3.5527e-15 | -0.0013759 | 0.0013759
| s2 = | 8.5714e-05 | 4.5701e-05 | 0.00021319
| s = | 0.0092582 | 0.0067602 | 0.014601
| df = | 14
| t = | 2.1448
Comparing quadratic and cubic models:
| F ratio = | 0.93333
| F critical = | 2.463 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The quadratic model appears to be a better choice than cubic.
| Quartic Model |  | Expanded Uncertainty
| parameter | estimate | lower | upper
| b0 = | 0.2 | 0.18804 | 0.21196
| b1 = | 1 | 0.9559 | 1.0441
| b2 = | -0.05 | -0.092129 | -0.0078707
| b3 = | 1.7053e-13 | -0.013294 | 0.013294
| b4 = | -2.1316e-14 | -0.0013211 | 0.0013211
| s2 = | 9.2308e-05 | 4.8249e-05 | 0.00023958
| s = | 0.0096077 | 0.0069461 | 0.015478
| df = | 13
| t = | 2.1604
Comparing quadratic and cubic models:
| F ratio = | 0.92857
| F critical = | 2.5536 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The cubic model appears to be a better choice than quartic.
| STEP 3 |  | Review the Charts
View the charts named Curves 1 and Residuals 1. Curves 1 shows bands
illustrating expanded uncertainties for the estimated regression. Compare these
bands with those of the quadratic regression, Curves 2.
(Note: You can change the associated approximate confidence level by
changing the "p-value" in cell F4 of the worksheet named Chart Data.)
Residuals 1 shows how the calibration points deviated from the calibration
line. Look for a simple pattern (such as a quadratic curve) in the chart. If
such a pattern appears, the quadratic model may be better. View Residual 2,
the deviations from the best- fit quadratic curve. If Residual 2 effectively
removes the simple pattern observed in Residual 1 and if the magnitude of
the deviations has been significantly reduced (as evidenced by a reduction
in the uncertainty u2), then the quadratic model is superior. An F-test can be
run to determine if the two overall uncertainties of fit are significantly different.
| F = | 219.6875 |  | Prob. of greater F = | 7.5191e-15
The quadratic model produces a significantly smaller overall uncertainty of fit.
The quadratic model appears to be the better choice.
| STEP 4 |  | Test for Zero Material on Same Day as Initial Calibration | [8 empty] | STEP 4 |  | Test for Zero Material on Same Day as Initial Calibration |  |  | STEP 4 |  | Test for Zero Material on Same Day as Initial Calibration | [13 empty] | 1: STEP 4; 2: Assess Magnitude of Uncertainty
This test can only be used when using the linear or quadratic model.
| Enter the impurity specification for the zero air material estimation and prediction | [10 empty] | Calculations - Linear Model |  |  |  |  | Calculations - Quadratic Model | [15 empty] | 1: Calculations - Quartic Model
[33 empty] | x | xb | x'V | x'Vx | t(0.9,df)
|  | Maximum Concentration = | 0.1 | [8 empty] | x | xb | x'V | x'Vx | t(0.9,df) | x | xb | x'V | x'Vx | t(0.9,df) | [11 empty] | 1 | 0.2995 | 2.5374e-05 | 2.3565e-05 | 1.7709
[12 empty] | 1 | 0.44167 | 0.002985 | 0.0029046 | 1.7459 | 1 | 0.2995 | 2.0357e-05 | 1.8996e-05 | 1.7531 | [11 empty] | 1: 0.1; 2: -1.7987e-05
| Maximum Response that can be measured at 95% confidence to make sure concentration is less than | [10 empty] | 0.1 |  | -0.00080343 |  |  | 0.1 |  | -1.3812e-05 | [13 empty] | 1: 0.01; 2: -1.3001e-06
| Maximum Concentration when Predicting Response. | [15 empty] | 0.01 |  | 2.031e-06 | [13 empty] | 1: 0.001; 2: 2.3529e-06
[33 empty] | 0.0001 |  | -3.162e-07
0.28206
[12 empty] | 90% Prediction Limits for Response.(for one tailed test) |  |  |  |  | 90% Prediction Limits for Response.(for one tailed test)
[12 empty] | Estimate | Lower | Upper |  |  | Estimate | Lower | Upper
[12 empty] | 0.44167 | 0.19182 | 0.69151 |  |  | 0.2995 | 0.28206 | 0.31694
| STEP 5 |  | Test for Zero Material on Different Day
from Initial Calibration
| This step applies to zero material tests on a different day than the initial calibration. | [10 empty] | STEP 5 |  | Test for Zero Air Material on Different Day |  |  | STEP 5 |  | Test for Zero Air Material on Different Day
| This test can only be used when using the linear or quadratic model. Before the zero material limit is input, | [12 empty] | from Initial Calibration |  |  |  |  | from Initial Calibration
| the measurement system is challenged with zero and span checks. Three or | [10 empty] | X'X |  |  |  |  | X'X
| more discrete checks of the zero gas and three or more checks of each span | [7 empty] | Conc | Conc^2 | Response | 18 | 45 | 0 | 0 |  | 18 | 45 | 165 | 0 | 0 | 0
| gas are made. If necessary, second span should be higher concentration than the first. | [7 empty] | 0 | 0 | 0.2 | 45 | 165 | 0 | 0 |  | 45 | 165 | 675 | 0 | 0 | 0
| Enter the results below: | [7 empty] | 0 | 0 | 0.21 | 0 | 0 | 9 | 9 |  | 165 | 675 | 2937 | 0 | 0 | 0
| 2 different span values need to be tested in the quadratic model | [7 empty] | 0 | 0 | 0.22 | 0 | 0 | 9 | 15 |  | 0 | 0 | 0 | 9 | 9 | 15
[7 empty] | (Only Quadratic Models) | [9 empty] | 0 | 0 | 0 | 9 | 15 | 27
|  |  |  | First |  |  | Second |  | 1 | 1 | 1.15 | [5 empty] | 0 | 0 | 0 | 15 | 27 | 51
|  | Response to |  | Span | Response |  | Span | Response | 1 | 1 | 1.16 | Det(X'X) | UpLeft | DnRight
|  | Zero gas |  | Conc. | to Span |  | Conc. | to Span | 1 | 1 | 1.17 | 51030 | 945 | 54 |  |  | Det(X'X) | UpLeft | DnRight
|  | 0.2 |  | 1 | 1.15 |  | 2 | 2 | [8 empty] | 11430720 | 105840 | 108
|  | 0.21 |  | 1 | 1.16 |  | 2 | 2.01 | 2 | 4 | 2 | X'Xinverse
|  | 0.22 |  | 1 | 1.17 |  | 2 | 2.02 | 2 | 4 | 2.01 | 0.1746 | -0.047619 | 0 | 0 |  | X'Xinverse
[9 empty] | 2 | 4 | 2.02 | -0.047619 | 0.019048 | 0 | 0 |  | 0.27381 | -0.19643 | 0.029762 | 0 | 0 | 0
|  |  |  | Max. Span Calculations | [7 empty] | 0 | 0 | 0.27778 | -0.16667 |  | -0.19643 | 0.24226 | -0.044643 | 0 | 0 | 0
| n = | 3 |  | n = | 3 | [6 empty] | 0 | 0 | -0.16667 | 0.16667 |  | 0.029762 | -0.044643 | 0.0089286 | 0 | 0 | 0
| mean = | 0.21 |  | mean = | 2.01 |  |  |  | nn = |  | 9 | [5 empty] | 0 | 0 | 0 | 0.33333 | -0.5 | 0.16667
| s = | 0.01 |  | s = | 0.01 | s = experimental standard deviation | [10 empty] | 0 | 0 | 0 | -0.5 | 2.1667 | -1
| Cal. Resp. = | 0.2 |  | Cal. Response = | 2 | [6 empty] | X'Y | b | Y'Y | df |  | 0 | 0 | 0 | 0.16667 | -1 | 0.5
[12 empty] | 40.35 | 0.36667 | 136.5537 | 23
|  |  | Zero Gas Results |  | Max. Span Gas Results | [6 empty] | 140.25 | 0.75 |  |  |  | X'Y | b | Y'Y | df
|  | Std. Uncertainty = s/sqrt(n) = | 0.0057735 |  | 0.0057735 | [6 empty] | 10.14 | 0.22667 | t | b'X'Y |  | 40.35 | 0.2 | 136.5537 | 21
|  | Rrs/100 = | 0.0201 |  | 0.0201 | [6 empty] | 15.54 | 0.9 | 2.0687 | 136.2669 |  | 140.25 | 1
|  |  | Std. Uncertainty is okay. |  | Std. Uncertainy is okay. | [11 empty] | 561.15 | -0.05
|  | Relative Difference (RD)= | -0.5% |  | 0.5% | [6 empty] | SSregr | SSresid | s2 | s |  | 10.14 | 0.21 | t | b'X'Y
|  |  | RD is okay. |  | RD is okay. | [6 empty] | 41.8506 | 0.2868 | 0.01247 | 0.11167 |  | 15.54 | 1 | 2.0796 | 136.5519
[17 empty] | 27.6 | -0.05
RD calculations above are based on linear/quadratic model.
| Following successful completion of the zero and span checks, zero material | [15 empty] | SSregr | SSresid | s2 | s
| limit can be predicted by the appropriate model | [15 empty] | 42.1356 | 0.0018 | 8.5714e-05 | 0.0092582
Coefficients are not significantly different.
Consider including the new data as part of original calibration (Step 1).
[12 empty] | STEP 5 |  | Test for Zero Air Material on Different Day |  |  | STEP 5 |  | Test for Zero Air Material on Different Day
[14 empty] | from Initial Calibration |  |  |  |  | from Initial Calibration
| STEP 5 | (Continued) | [9 empty] | Calculations - Linear Model |  |  |  |  | Calculations - Quadratic Model
|  | Maximum Concentration = | 0.1 | [8 empty] | Var-Covar |  |  |  |  | Var-Covar
[12 empty] | 0.0021772 | -0.00059379 | 0 | 0 |  | 2.3469e-05 | -1.6837e-05 | 2.551e-06 | 0 | 0 | 0
| Maximum Response that can be measured at 95% confidence to make sure concentration is less than | [10 empty] | -0.00059379 | 0.00023752 | 0 | 0 |  | -1.6837e-05 | 2.0765e-05 | -3.8265e-06 | 0 | 0 | 0
| Maximum Concentration when Predicting. | [10 empty] | 0 | 0 | 0.0034638 | -0.0020783 |  | 2.551e-06 | -3.8265e-06 | 7.6531e-07 | 0 | 0 | 0
[12 empty] | 0 | 0 | -0.0020783 | 0.0020783 |  | 0 | 0 | 0 | 2.8571e-05 | -4.2857e-05 | 1.4286e-05
|  |  | 0.29164 | [13 empty] | 0 | 0 | 0 | -4.2857e-05 | 0.00018571 | -8.5714e-05
[12 empty] | Have Coefficients Remained the Same? |  |  |  |  | 0 | 0 | 0 | 1.4286e-05 | -8.5714e-05 | 4.2857e-05
[12 empty] | b0-b2 | -0.01537 | 0.29537 | TRUE
[12 empty] | b1-b3 | -0.24955 | -0.050451 | FALSE |  | Have Coefficients Remained the Same?
[17 empty] | b0-b3 | -0.01 | -0.025002 | 0.0050022 | TRUE
[12 empty] | One coefficient is significantly different, indicating a calibration drift. |  |  |  |  | b1-b4 | -3.5527e-15 | -0.029883 | 0.029883 | TRUE
[12 empty] | Estimated concentration includes correction for calibration drift. |  |  |  |  | b2-b5 | -8.8818e-16 | -0.013735 | 0.013735 | TRUE
[12 empty] | x | xb | x'V | x'Vx | t(0.9,df) | Coefficients are not significantly different.
[12 empty] | 0 | 0.31667 | 0 | 0.0030689 | 1.7139 | Consider including the new data as part of original calibration (Step 1).
[12 empty] | 0 |  | 0
[12 empty] | 1 |  | 0.0032559 |  |  | x | xb | x'V | x'Vx | t(0.9,df)
[12 empty] | 0.1 |  | -0.0018704 |  |  | 0 | 0.3095 | 0 | 2.1976e-05 | 1.7207
[17 empty] | 0 |  | 0
[17 empty] | 0 |  | 0
[17 empty] | 1 |  | 2.4429e-05
[17 empty] | 0.1 |  | -2.5143e-05
[12 empty] | 90% Prediction Limits for Response.(for one tailed test) |  |  |  |  | 0.01 |  | 6.1429e-06
[12 empty] | Estimate | Lower | Upper
[12 empty] | 0.31667 | 0.10303 | 0.65531
90% Prediction Limits for Response.(for one tailed test)
[17 empty] | Estimate | Lower | Upper
[17 empty] | 0.3095 | 0.29164 | 0.32736
Note on Z53: This compares the residuals of the linear and quadratic regressions. If significant, then quadratic is the preferred model.
Note on AI53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on BL53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on Z56: Comparing the Linear and Quadratic solutions. Q --> quadratic is better L --> linear is better model
Note on AI56: Comparing the Quadratic and Cubic solutions
Note on BL56: Comparing the Cubic and Quartic solutions

## Chart Data
This sheet contains the chart data. The chart data include initial calibration data and the
regression information derived from those data.
Only the initial calibration data (Step 1 in Measurement Data worksheet) are included.
| Recalibration data (Step 6) are not included. | [5 empty] | 0.95 | = Confidence Level for Plotting Confidence Bands
0 | = minimum concentration | [5 empty] | 2.1199 | = t for linear model
5 | = maximum concentration | [5 empty] | 2.1314 | = t for quadratic model
18 | = n, the number of initial calibration points | [5 empty] | 2.1448 | = t for cubic model
[7 empty] | 2.1604 | = t for quartic model
Initial Calibration Data |  |  |  |  | Estimates, Linear Model |  |  |  | Estimates, Quadratic Model |  |  |  | Estimates, Cubic Model |  |  |  | Estimates, Quartic Model
i | Xi | Yi |  |  | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual
1 | 0 | 0.19 | [5 empty] | -0.17667 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
2 | 0 | 0.2 | [5 empty] | -0.16667 |  |  |  | -2.8311e-15 |  |  |  | 1.138e-14 |  |  |  | 8.2434e-14
3 | 0 | 0.21 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
4 | 1 | 1.14 | [5 empty] | 0.023333 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
5 | 1 | 1.15 | [5 empty] | 0.033333 |  |  |  | -3.1086e-15 |  |  |  | -1.088e-13 |  |  |  | 1.1502e-13
6 | 1 | 1.16 | [5 empty] | 0.043333 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
7 | 2 | 1.99 | [5 empty] | 0.12333 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
8 | 2 | 2 | [5 empty] | 0.13333 |  |  |  | -3.5527e-15 |  |  |  | -2.558e-13 |  |  |  | -2.1316e-13
9 | 2 | 2.01 | [5 empty] | 0.14333 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
10 | 3 | 2.74 | [5 empty] | 0.12333 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
11 | 3 | 2.75 | [5 empty] | 0.13333 |  |  |  | -4.4409e-15 |  |  |  | -4.5119e-13 |  |  |  | -1.1582e-12
12 | 3 | 2.76 | [5 empty] | 0.14333 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
13 | 4 | 3.39 | [5 empty] | 0.023333 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
14 | 4 | 3.4 | [5 empty] | 0.033333 |  |  |  | -5.7732e-15 |  |  |  | -7.1632e-13 |  |  |  | -2.4643e-12
15 | 4 | 3.41 | [5 empty] | 0.043333 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
16 | 5 | 3.94 | [5 empty] | -0.17667 |  |  |  | -0.01 |  |  |  | -0.01 |  |  |  | -0.01
17 | 5 | 3.95 | [5 empty] | -0.16667 |  |  |  | -7.1054e-15 |  |  |  | -1.072e-12 |  |  |  | -3.3635e-12
18 | 5 | 3.96 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
19 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
20 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
21 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
22 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
23 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
24 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
25 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
26 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
27 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
28 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
29 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
30 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
31 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
32 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
33 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
34 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
35 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
36 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
37 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
38 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
39 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
40 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
41 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
42 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
43 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
44 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
45 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
46 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
47 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
48 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
49 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
50 | -999 | -999 | [5 empty] | -0.15667 |  |  |  | 0.01 |  |  |  | 0.01 |  |  |  | 0.01
[21 empty] | x'V for Linear |  | x'V for Quadratic |  |  | x'V for Cubic |  |  |  | x'V for Quartic
1 | 0 | 0 | 0 | 0 | 0.36667 | 0.24923 | 0.4841 |  | 0.2 | 0.19002 | 0.20998 |  | 0.2 | 0.19259 | 0.20741 |  | 0.2 | 0.19186 | 0.20814 |  | 0.0030687 | -0.0008369 | 2.1905e-05 | -1.5714e-05 | 2.381e-06 | 2.5608e-05 | -3.2628e-05 | 1.164e-05 | -1.2346e-06 | 3.0647e-05 | -5.5963e-05 | 3.312e-05 | -7.8348e-06 | 6.4103e-07
1 | 0.1 | 0.01 | 0.001 | 0.0001 | 0.44167 | 0.32742 | 0.55592 |  | 0.2995 | 0.29021 | 0.30879 |  | 0.2995 | 0.29292 | 0.30608 |  | 0.2995 | 0.29236 | 0.30664 |  | 0.002985 | -0.00080343 | 2.0357e-05 | -1.3812e-05 | 2.031e-06 | 2.2461e-05 | -2.3419e-05 | 7.2902e-06 | -7.0123e-07 | 2.5374e-05 | -1.7987e-05 | -1.3001e-06 | 2.3529e-06 | -3.162e-07
1 | 0.2 | 0.04 | 0.008 | 0.0016 | 0.51667 | 0.40555 | 0.62778 |  | 0.398 | 0.38933 | 0.40667 |  | 0.398 | 0.39202 | 0.40398 |  | 0.398 | 0.39107 | 0.40493 |  | 0.0029013 | -0.00076995 | 1.8857e-05 | -1.1981e-05 | 1.6952e-06 | 1.9539e-05 | -1.5093e-05 | 3.3989e-06 | -2.2716e-07 | 2.0718e-05 | 1.3032e-05 | -2.8809e-05 | 1.0394e-05 | -1.0656e-06
1 | 0.3 | 0.09 | 0.027 | 0.0081 | 0.59167 | 0.48364 | 0.69969 |  | 0.4955 | 0.48739 | 0.50361 |  | 0.4955 | 0.48991 | 0.50109 |  | 0.4955 | 0.48832 | 0.50268 |  | 0.0028176 | -0.00073648 | 1.7405e-05 | -1.0221e-05 | 1.3738e-06 | 1.6834e-05 | -7.6167e-06 | -5.2116e-08 | 1.9012e-07 | 1.6633e-05 | 3.7737e-05 | -5.0075e-05 | 1.6503e-05 | -1.6284e-06
1 | 0.4 | 0.16 | 0.064 | 0.0256 | 0.66667 | 0.56168 | 0.77165 |  | 0.592 | 0.58439 | 0.59961 |  | 0.592 | 0.58662 | 0.59738 |  | 0.592 | 0.58443 | 0.59957 |  | 0.0027339 | -0.000703 | 1.6e-05 | -8.5333e-06 | 1.0667e-06 | 1.4341e-05 | -9.5605e-07 | -3.0815e-06 | 5.5309e-07 | 1.3076e-05 | 5.6743e-05 | -6.574e-05 | 2.0884e-05 | -2.0246e-06
1 | 0.5 | 0.25 | 0.125 | 0.0625 | 0.74167 | 0.63966 | 0.84367 |  | 0.6875 | 0.68031 | 0.69469 |  | 0.6875 | 0.68218 | 0.69282 |  | 0.6875 | 0.67958 | 0.69542 |  | 0.0026502 | -0.00066952 | 1.4643e-05 | -6.9167e-06 | 7.7381e-07 | 1.205e-05 | 4.9228e-06 | -5.7077e-06 | 8.642e-07 | 1.0006e-05 | 7.0641e-05 | -7.6419e-05 | 2.3734e-05 | -2.2736e-06
1 | 0.6 | 0.36 | 0.216 | 0.1296 | 0.81667 | 0.71758 | 0.91576 |  | 0.782 | 0.77516 | 0.78884 |  | 0.782 | 0.77665 | 0.78735 |  | 0.782 | 0.77385 | 0.79015 |  | 0.0025665 | -0.00063605 | 1.3333e-05 | -5.3714e-06 | 4.9524e-07 | 9.9556e-06 | 1.0054e-05 | -7.9492e-06 | 1.1259e-06 | 7.3834e-06 | 7.9996e-05 | -8.2697e-05 | 2.5238e-05 | -2.3938e-06
1 | 0.7 | 0.49 | 0.343 | 0.2401 | 0.89167 | 0.79543 | 0.98791 |  | 0.8755 | 0.86895 | 0.88205 |  | 0.8755 | 0.87005 | 0.88095 |  | 0.8755 | 0.86727 | 0.88373 |  | 0.0024828 | -0.00060257 | 1.2071e-05 | -3.8976e-06 | 2.3095e-07 | 8.0492e-06 | 1.4471e-05 | -9.8246e-06 | 1.3407e-06 | 5.1686e-06 | 8.5346e-05 | -8.5135e-05 | 2.5574e-05 | -2.4027e-06
1 | 0.8 | 0.64 | 0.512 | 0.4096 | 0.96667 | 0.8732 | 1.0601 |  | 0.968 | 0.96166 | 0.97434 |  | 0.968 | 0.96244 | 0.97356 |  | 0.968 | 0.95982 | 0.97618 |  | 0.0023991 | -0.0005691 | 1.0857e-05 | -2.4952e-06 | -1.9048e-08 | 6.3238e-06 | 1.8207e-05 | -1.1352e-05 | 1.5111e-06 | 3.3248e-06 | 8.7206e-05 | -8.4262e-05 | 2.4913e-05 | -2.3169e-06
1 | 0.9 | 0.81 | 0.729 | 0.6561 | 1.0417 | 0.9509 | 1.1324 |  | 1.0595 | 1.0533 | 1.0657 |  | 1.0595 | 1.0538 | 1.0652 |  | 1.0595 | 1.0515 | 1.0675 |  | 0.0023154 | -0.00053562 | 9.6905e-06 | -1.1643e-06 | -2.5476e-07 | 4.772e-06 | 2.1297e-05 | -1.2551e-05 | 1.6395e-06 | 1.8168e-06 | 8.6062e-05 | -8.0582e-05 | 2.3413e-05 | -2.1521e-06
1 | 1 | 1 | 1 | 1 | 1.1167 | 1.0285 | 1.2048 |  | 1.15 | 1.1439 | 1.1561 |  | 1.15 | 1.1442 | 1.1558 |  | 1.15 | 1.1423 | 1.1577 |  | 0.0022317 | -0.00050214 | 8.5714e-06 | 9.5238e-08 | -4.7619e-07 | 3.3862e-06 | 2.3774e-05 | -1.3439e-05 | 1.7284e-06 | 6.105e-07 | 8.2377e-05 | -7.4573e-05 | 2.1225e-05 | -1.9231e-06
1 | 1.1 | 1.21 | 1.331 | 1.4641 | 1.1917 | 1.106 | 1.2773 |  | 1.2395 | 1.2334 | 1.2456 |  | 1.2395 | 1.2337 | 1.2453 |  | 1.2395 | 1.2321 | 1.2469 |  | 0.0021481 | -0.00046867 | 7.5e-06 | 1.2833e-06 | -6.8333e-07 | 2.1593e-06 | 2.5673e-05 | -1.4035e-05 | 1.7802e-06 | -3.2643e-07 | 7.6587e-05 | -6.6682e-05 | 1.8492e-05 | -1.6438e-06
1 | 1.2 | 1.44 | 1.728 | 2.0736 | 1.2667 | 1.1834 | 1.3499 |  | 1.328 | 1.3219 | 1.3341 |  | 1.328 | 1.3222 | 1.3338 |  | 1.328 | 1.3209 | 1.3351 |  | 0.0020644 | -0.00043519 | 6.4762e-06 | 2.4e-06 | -8.7619e-07 | 1.0836e-06 | 2.7026e-05 | -1.4358e-05 | 1.7975e-06 | -1.0249e-06 | 6.9103e-05 | -5.733e-05 | 1.5346e-05 | -1.3272e-06
1 | 1.3 | 1.69 | 2.197 | 2.8561 | 1.3417 | 1.2607 | 1.4226 |  | 1.4155 | 1.4094 | 1.4216 |  | 1.4155 | 1.4097 | 1.4213 |  | 1.4155 | 1.4088 | 1.4222 |  | 0.0019807 | -0.00040171 | 5.5e-06 | 3.4452e-06 | -1.0548e-06 | 1.5185e-07 | 2.7868e-05 | -1.4425e-05 | 1.7827e-06 | -1.5141e-06 | 6.0311e-05 | -4.6911e-05 | 1.1911e-05 | -9.8543e-07
1 | 1.4 | 1.96 | 2.744 | 3.8416 | 1.4167 | 1.3379 | 1.4955 |  | 1.502 | 1.4958 | 1.5082 |  | 1.502 | 1.4962 | 1.5078 |  | 1.502 | 1.4956 | 1.5084 |  | 0.001897 | -0.00036824 | 4.5714e-06 | 4.419e-06 | -1.219e-06 | -6.4339e-07 | 2.8233e-05 | -1.4256e-05 | 1.7383e-06 | -1.8219e-06 | 5.057e-05 | -3.5791e-05 | 8.3031e-06 | -6.2974e-07
1 | 1.5 | 2.25 | 3.375 | 5.0625 | 1.4917 | 1.4149 | 1.5684 |  | 1.5875 | 1.5813 | 1.5937 |  | 1.5875 | 1.5818 | 1.5932 |  | 1.5875 | 1.5813 | 1.5937 |  | 0.0018133 | -0.00033476 | 3.6905e-06 | 5.3214e-06 | -1.369e-06 | -1.3095e-06 | 2.8155e-05 | -1.3869e-05 | 1.6667e-06 | -1.9746e-06 | 4.0213e-05 | -2.4309e-05 | 4.6274e-06 | -2.7043e-07
1 | 1.6 | 2.56 | 4.096 | 6.5536 | 1.5667 | 1.4918 | 1.6415 |  | 1.672 | 1.6657 | 1.6783 |  | 1.672 | 1.6664 | 1.6776 |  | 1.672 | 1.6659 | 1.6781 |  | 0.0017296 | -0.00030129 | 2.8571e-06 | 6.1524e-06 | -1.5048e-06 | -1.854e-06 | 2.7666e-05 | -1.3283e-05 | 1.5704e-06 | -1.9968e-06 | 2.9549e-05 | -1.2774e-05 | 9.812e-07 | 8.3077e-08
1 | 1.7 | 2.89 | 4.913 | 8.3521 | 1.6417 | 1.5685 | 1.7148 |  | 1.7555 | 1.7491 | 1.7619 |  | 1.7555 | 1.75 | 1.761 |  | 1.7555 | 1.7495 | 1.7615 |  | 0.0016459 | -0.00026781 | 2.0714e-06 | 6.9119e-06 | -1.6262e-06 | -2.2841e-06 | 2.6802e-05 | -1.2515e-05 | 1.4519e-06 | -1.9117e-06 | 1.8861e-05 | -1.4711e-06 | -2.5474e-06 | 4.2226e-07
1 | 1.8 | 3.24 | 5.832 | 10.4976 | 1.7167 | 1.6451 | 1.7883 |  | 1.838 | 1.8315 | 1.8445 |  | 1.838 | 1.8326 | 1.8434 |  | 1.838 | 1.8319 | 1.8441 |  | 0.0015622 | -0.00023433 | 1.3333e-06 | 7.6e-06 | -1.7333e-06 | -2.6074e-06 | 2.5596e-05 | -1.1585e-05 | 1.3136e-06 | -1.7409e-06 | 8.4057e-06 | 9.3453e-06 | -5.8792e-06 | 7.3949e-07
1 | 1.9 | 3.61 | 6.859 | 13.0321 | 1.7917 | 1.7215 | 1.8619 |  | 1.9195 | 1.9129 | 1.9261 |  | 1.9195 | 1.9142 | 1.9248 |  | 1.9195 | 1.9133 | 1.9257 |  | 0.0014785 | -0.00020086 | 6.4286e-07 | 8.2167e-06 | -1.8262e-06 | -2.8312e-06 | 2.4082e-05 | -1.0511e-05 | 1.158e-06 | -1.5045e-06 | -1.5858e-06 | 1.9447e-05 | -8.9441e-06 | 1.028e-06
1 | 2 | 4 | 8 | 16 | 1.8667 | 1.7976 | 1.9357 |  | 2 | 1.9933 | 2.0067 |  | 2 | 1.9949 | 2.0051 |  | 2 | 1.9937 | 2.0063 |  | 0.0013948 | -0.00016738 | -3.3881e-21 | 8.7619e-06 | -1.9048e-06 | -2.963e-06 | 2.2293e-05 | -9.3122e-06 | 9.8765e-07 | -1.221e-06 | -1.0908e-05 | 2.8632e-05 | -1.1681e-05 | 1.2821e-06
1 | 2.1 | 4.41 | 9.261 | 19.4481 | 1.9417 | 1.8736 | 2.0097 |  | 2.0795 | 2.0727 | 2.0863 |  | 2.0795 | 2.0745 | 2.0845 |  | 2.0795 | 2.073 | 2.086 |  | 0.0013112 | -0.0001339 | -5.9524e-07 | 9.2357e-06 | -1.969e-06 | -3.0101e-06 | 2.0263e-05 | -8.0061e-06 | 8.0494e-07 | -9.0751e-07 | -1.938e-05 | 3.673e-05 | -1.4037e-05 | 1.4966e-06
1 | 2.2 | 4.84 | 10.648 | 23.4256 | 2.0167 | 1.9494 | 2.0839 |  | 2.158 | 2.1512 | 2.1648 |  | 2.158 | 2.1531 | 2.1629 |  | 2.158 | 2.1514 | 2.1646 |  | 0.0012275 | -0.00010043 | -1.1429e-06 | 9.6381e-06 | -2.019e-06 | -2.9799e-06 | 1.8027e-05 | -6.6116e-06 | 6.1235e-07 | -5.7944e-07 | -2.6848e-05 | 4.3593e-05 | -1.597e-05 | 1.6677e-06
1 | 2.3 | 5.29 | 12.167 | 27.9841 | 2.0917 | 2.025 | 2.1584 |  | 2.2355 | 2.2286 | 2.2424 |  | 2.2355 | 2.2307 | 2.2403 |  | 2.2355 | 2.2288 | 2.2422 |  | 0.0011438 | -6.6952e-05 | -1.6429e-06 | 9.969e-06 | -2.0548e-06 | -2.8799e-06 | 1.5618e-05 | -5.1474e-06 | 4.1235e-07 | -2.5073e-07 | -3.3183e-05 | 4.9105e-05 | -1.7446e-05 | 1.7921e-06
1 | 2.4 | 5.76 | 13.824 | 33.1776 | 2.1667 | 2.1003 | 2.233 |  | 2.312 | 2.3051 | 2.3189 |  | 2.312 | 2.3072 | 2.3168 |  | 2.312 | 2.3052 | 2.3188 |  | 0.0010601 | -3.3476e-05 | -2.0952e-06 | 1.0229e-05 | -2.0762e-06 | -2.7175e-06 | 1.307e-05 | -3.6317e-06 | 2.0741e-07 | 6.6227e-08 | -3.8282e-05 | 5.3174e-05 | -1.8438e-05 | 1.8677e-06
1 | 2.5 | 6.25 | 15.625 | 39.0625 | 2.2417 | 2.1754 | 2.3079 |  | 2.3875 | 2.3806 | 2.3944 |  | 2.3875 | 2.3827 | 2.3923 |  | 2.3875 | 2.3806 | 2.3944 |  | 0.00097639 | 4.3368e-19 | -2.5e-06 | 1.0417e-05 | -2.0833e-06 | -2.5e-06 | 1.0417e-05 | -2.0833e-06 | 2.7952e-20 | 3.6058e-07 | -4.2067e-05 | 5.5739e-05 | -1.893e-05 | 1.893e-06
1 | 2.6 | 6.76 | 17.576 | 45.6976 | 2.3167 | 2.2503 | 2.383 |  | 2.462 | 2.4551 | 2.4689 |  | 2.462 | 2.4572 | 2.4668 |  | 2.462 | 2.4552 | 2.4688 |  | 0.0008927 | 3.3476e-05 | -2.8571e-06 | 1.0533e-05 | -2.0762e-06 | -2.2349e-06 | 7.6919e-06 | -5.2063e-07 | -2.0741e-07 | 6.23e-07 | -4.4487e-05 | 5.6764e-05 | -1.8916e-05 | 1.8677e-06
1 | 2.7 | 7.29 | 19.683 | 53.1441 | 2.3917 | 2.325 | 2.4584 |  | 2.5355 | 2.5286 | 2.5424 |  | 2.5355 | 2.5307 | 2.5403 |  | 2.5355 | 2.5288 | 2.5422 |  | 0.00080901 | 6.6952e-05 | -3.1667e-06 | 1.0579e-05 | -2.0548e-06 | -1.9296e-06 | 4.9294e-06 | 1.0378e-06 | -4.1235e-07 | 8.4573e-07 | -4.5516e-05 | 5.6242e-05 | -1.8397e-05 | 1.7921e-06
1 | 2.8 | 7.84 | 21.952 | 61.4656 | 2.4667 | 2.3994 | 2.5339 |  | 2.608 | 2.6012 | 2.6148 |  | 2.608 | 2.6031 | 2.6129 |  | 2.608 | 2.6014 | 2.6146 |  | 0.00072532 | 0.00010043 | -3.4286e-06 | 1.0552e-05 | -2.019e-06 | -1.5915e-06 | 2.1632e-06 | 2.5735e-06 | -6.1235e-07 | 1.0225e-06 | -4.5152e-05 | 5.4191e-05 | -1.7383e-05 | 1.6677e-06
1 | 2.9 | 8.41 | 24.389 | 70.7281 | 2.5417 | 2.4736 | 2.6097 |  | 2.6795 | 2.6727 | 2.6863 |  | 2.6795 | 2.6745 | 2.6845 |  | 2.6795 | 2.673 | 2.686 |  | 0.00064163 | 0.0001339 | -3.6429e-06 | 1.0455e-05 | -1.969e-06 | -1.228e-06 | -5.7289e-07 | 4.068e-06 | -8.0494e-07 | 1.1487e-06 | -4.3422e-05 | 5.0661e-05 | -1.5895e-05 | 1.4966e-06
1 | 3 | 9 | 27 | 81 | 2.6167 | 2.5476 | 2.6857 |  | 2.75 | 2.7433 | 2.7567 |  | 2.75 | 2.7449 | 2.7551 |  | 2.75 | 2.7437 | 2.7563 |  | 0.00055794 | 0.00016738 | -3.8095e-06 | 1.0286e-05 | -1.9048e-06 | -8.4656e-07 | -3.2451e-06 | 5.5026e-06 | -9.8765e-07 | 1.221e-06 | -4.0374e-05 | 4.5726e-05 | -1.396e-05 | 1.2821e-06
1 | 3.1 | 9.61 | 29.791 | 92.3521 | 2.6917 | 2.6215 | 2.7619 |  | 2.8195 | 2.8129 | 2.8261 |  | 2.8195 | 2.8142 | 2.8248 |  | 2.8195 | 2.8133 | 2.8257 |  | 0.00047425 | 0.00020086 | -3.9286e-06 | 1.0045e-05 | -1.8262e-06 | -4.545e-07 | -5.8197e-06 | 6.859e-06 | -1.158e-06 | 1.2379e-06 | -3.6087e-05 | 3.9489e-05 | -1.1616e-05 | 1.028e-06
1 | 3.2 | 10.24 | 32.768 | 104.8576 | 2.7667 | 2.6951 | 2.8383 |  | 2.888 | 2.8815 | 2.8945 |  | 2.888 | 2.8826 | 2.8934 |  | 2.888 | 2.8819 | 2.8941 |  | 0.00039056 | 0.00023433 | -4e-06 | 9.7333e-06 | -1.7333e-06 | -5.9259e-08 | -8.2627e-06 | 8.1185e-06 | -1.3136e-06 | 1.1993e-06 | -3.0662e-05 | 3.208e-05 | -8.9105e-06 | 7.3949e-07
1 | 3.3 | 10.89 | 35.937 | 118.5921 | 2.8417 | 2.7685 | 2.9148 |  | 2.9555 | 2.9491 | 2.9619 |  | 2.9555 | 2.95 | 2.961 |  | 2.9555 | 2.9495 | 2.9615 |  | 0.00030687 | 0.00026781 | -4.0238e-06 | 9.35e-06 | -1.6262e-06 | 3.3175e-07 | -1.054e-05 | 9.2627e-06 | -1.4519e-06 | 1.1067e-06 | -2.4227e-05 | 2.3657e-05 | -5.8978e-06 | 4.2226e-07
1 | 3.4 | 11.56 | 39.304 | 133.6336 | 2.9167 | 2.8418 | 2.9915 |  | 3.022 | 3.0157 | 3.0283 |  | 3.022 | 3.0164 | 3.0276 |  | 3.022 | 3.0159 | 3.0281 |  | 0.00022317 | 0.00030129 | -4e-06 | 8.8952e-06 | -1.5048e-06 | 7.1111e-07 | -1.2619e-05 | 1.0273e-05 | -1.5704e-06 | 9.6293e-07 | -1.6934e-05 | 1.4405e-05 | -2.6427e-06 | 8.3077e-08
1 | 3.5 | 12.25 | 42.875 | 150.0625 | 2.9917 | 2.9149 | 3.0684 |  | 3.0875 | 3.0813 | 3.0937 |  | 3.0875 | 3.0818 | 3.0932 |  | 3.0875 | 3.0813 | 3.0937 |  | 0.00013948 | 0.00033476 | -3.9286e-06 | 8.369e-06 | -1.369e-06 | 1.0714e-06 | -1.4464e-05 | 1.1131e-05 | -1.6667e-06 | 7.7266e-07 | -8.9629e-06 | 4.5373e-06 | 7.8125e-07 | -2.7043e-07
1 | 3.6 | 12.96 | 46.656 | 167.9616 | 3.0667 | 2.9879 | 3.1455 |  | 3.152 | 3.1458 | 3.1582 |  | 3.152 | 3.1462 | 3.1578 |  | 3.152 | 3.1456 | 3.1584 |  | 5.5794e-05 | 0.00036824 | -3.8095e-06 | 7.7714e-06 | -1.219e-06 | 1.4053e-06 | -1.6043e-05 | 1.1818e-05 | -1.7383e-06 | 5.4193e-07 | -5.1836e-07 | -5.706e-06 | 4.2917e-06 | -6.2974e-07
1 | 3.7 | 13.69 | 50.653 | 187.4161 | 3.1417 | 3.0607 | 3.2226 |  | 3.2155 | 3.2094 | 3.2216 |  | 3.2155 | 3.2097 | 3.2213 |  | 3.2155 | 3.2088 | 3.2222 |  | -2.7897e-05 | 0.00040171 | -3.6429e-06 | 7.1024e-06 | -1.0548e-06 | 1.7053e-06 | -1.7321e-05 | 1.2316e-05 | -1.7827e-06 | 2.7833e-07 | 8.1696e-06 | -1.6057e-05 | 7.7973e-06 | -9.8543e-07
1 | 3.8 | 14.44 | 54.872 | 208.5136 | 3.2167 | 3.1334 | 3.2999 |  | 3.278 | 3.2719 | 3.2841 |  | 3.278 | 3.2722 | 3.2838 |  | 3.278 | 3.2709 | 3.2851 |  | -0.00011159 | 0.00043519 | -3.4286e-06 | 6.3619e-06 | -8.7619e-07 | 1.964e-06 | -1.8264e-05 | 1.2605e-05 | -1.7975e-06 | -8.9866e-09 | 1.6845e-05 | -2.6219e-05 | 1.1198e-05 | -1.3272e-06
1 | 3.9 | 15.21 | 59.319 | 231.3441 | 3.2917 | 3.206 | 3.3773 |  | 3.3395 | 3.3334 | 3.3456 |  | 3.3395 | 3.3337 | 3.3453 |  | 3.3395 | 3.3321 | 3.3469 |  | -0.00019528 | 0.00046867 | -3.1667e-06 | 5.55e-06 | -6.8333e-07 | 2.1741e-06 | -1.8839e-05 | 1.2669e-05 | -1.7802e-06 | -3.0934e-07 | 2.5227e-05 | -3.587e-05 | 1.4384e-05 | -1.6438e-06
1 | 4 | 16 | 64 | 256 | 3.3667 | 3.2785 | 3.4548 |  | 3.4 | 3.3939 | 3.4061 |  | 3.4 | 3.3942 | 3.4058 |  | 3.4 | 3.3923 | 3.4077 |  | -0.00027897 | 0.00050214 | -2.8571e-06 | 4.6667e-06 | -4.7619e-07 | 2.328e-06 | -1.9012e-05 | 1.2487e-05 | -1.7284e-06 | -6.105e-07 | 3.3008e-05 | -4.4658e-05 | 1.7236e-05 | -1.9231e-06
1 | 4.1 | 16.81 | 68.921 | 282.5761 | 3.4417 | 3.3509 | 3.5324 |  | 3.4595 | 3.4533 | 3.4657 |  | 3.4595 | 3.4538 | 3.4652 |  | 3.4595 | 3.4515 | 3.4675 |  | -0.00036266 | 0.00053562 | -2.5e-06 | 3.7119e-06 | -2.5476e-07 | 2.4185e-06 | -1.8749e-05 | 1.2042e-05 | -1.6395e-06 | -8.9871e-07 | 3.9855e-05 | -5.2206e-05 | 1.9629e-05 | -2.1521e-06
1 | 4.2 | 17.64 | 74.088 | 311.1696 | 3.5167 | 3.4232 | 3.6101 |  | 3.518 | 3.5117 | 3.5243 |  | 3.518 | 3.5124 | 3.5236 |  | 3.518 | 3.5098 | 3.5262 |  | -0.00044635 | 0.0005691 | -2.0952e-06 | 2.6857e-06 | -1.9048e-08 | 2.4381e-06 | -1.8017e-05 | 1.1314e-05 | -1.5111e-06 | -1.1587e-06 | 4.541e-05 | -5.8108e-05 | 2.1426e-05 | -2.3169e-06
1 | 4.3 | 18.49 | 79.507 | 341.8801 | 3.5917 | 3.4954 | 3.6879 |  | 3.5755 | 3.5689 | 3.5821 |  | 3.5755 | 3.5701 | 3.5809 |  | 3.5755 | 3.5673 | 3.5837 |  | -0.00053004 | 0.00060257 | -1.6429e-06 | 1.5881e-06 | 2.3095e-07 | 2.3794e-06 | -1.678e-05 | 1.0287e-05 | -1.3407e-06 | -1.3736e-06 | 4.9288e-05 | -6.1929e-05 | 2.248e-05 | -2.4027e-06
1 | 4.4 | 19.36 | 85.184 | 374.8096 | 3.6667 | 3.5676 | 3.7658 |  | 3.632 | 3.6252 | 3.6388 |  | 3.632 | 3.6266 | 3.6374 |  | 3.632 | 3.6238 | 3.6402 |  | -0.00061373 | 0.00063605 | -1.1429e-06 | 4.1905e-07 | 4.9524e-07 | 2.2349e-06 | -1.5006e-05 | 8.9397e-06 | -1.1259e-06 | -1.525e-06 | 5.1081e-05 | -6.321e-05 | 2.2639e-05 | -2.3938e-06
1 | 4.5 | 20.25 | 91.125 | 410.0625 | 3.7417 | 3.6397 | 3.8437 |  | 3.6875 | 3.6803 | 3.6947 |  | 3.6875 | 3.6822 | 3.6928 |  | 3.6875 | 3.6796 | 3.6954 |  | -0.00069742 | 0.00066952 | -5.9524e-07 | -8.2143e-07 | 7.7381e-07 | 1.9974e-06 | -1.2661e-05 | 7.2553e-06 | -8.642e-07 | -1.593e-06 | 5.0352e-05 | -6.1462e-05 | 2.1739e-05 | -2.2736e-06
1 | 4.6 | 21.16 | 97.336 | 447.7456 | 3.8167 | 3.7117 | 3.9217 |  | 3.742 | 3.7344 | 3.7496 |  | 3.742 | 3.7366 | 3.7474 |  | 3.742 | 3.7344 | 3.7496 |  | -0.00078111 | 0.000703 | -6.7763e-21 | -2.1333e-06 | 1.0667e-06 | 1.6593e-06 | -9.7106e-06 | 5.2148e-06 | -5.5309e-07 | -1.5562e-06 | 4.6642e-05 | -5.6168e-05 | 1.9608e-05 | -2.0246e-06
1 | 4.7 | 22.09 | 103.823 | 487.9681 | 3.8917 | 3.7836 | 3.9997 |  | 3.7955 | 3.7874 | 3.8036 |  | 3.7955 | 3.7899 | 3.8011 |  | 3.7955 | 3.7883 | 3.8027 |  | -0.0008648 | 0.00073648 | 6.4286e-07 | -3.5167e-06 | 1.3738e-06 | 1.2132e-06 | -6.1214e-06 | 2.7997e-06 | -1.9012e-07 | -1.3916e-06 | 3.9462e-05 | -4.6784e-05 | 1.6064e-05 | -1.6284e-06
1 | 4.8 | 23.04 | 110.592 | 530.8416 | 3.9667 | 3.8556 | 4.0778 |  | 3.848 | 3.8393 | 3.8567 |  | 3.848 | 3.842 | 3.854 |  | 3.848 | 3.8411 | 3.8549 |  | -0.00094849 | 0.00076995 | 1.3333e-06 | -4.9714e-06 | 1.6952e-06 | 6.5185e-07 | -1.8593e-06 | -8.4656e-09 | 2.2716e-07 | -1.0747e-06 | 2.8302e-05 | -3.274e-05 | 1.0919e-05 | -1.0656e-06
1 | 4.9 | 24.01 | 117.649 | 576.4801 | 4.0417 | 3.9274 | 4.1559 |  | 3.8995 | 3.8902 | 3.9088 |  | 3.8995 | 3.8929 | 3.9061 |  | 3.8995 | 3.8924 | 3.9066 |  | -0.0010322 | 0.00080343 | 2.0714e-06 | -6.4976e-06 | 2.031e-06 | -3.2275e-08 | 3.1093e-06 | -3.2283e-06 | 7.0123e-07 | -5.793e-07 | 1.2622e-05 | -1.3437e-05 | 3.9711e-06 | -3.162e-07
1 | 5 | 25 | 125 | 625 | 4.1167 | 3.9992 | 4.2341 |  | 3.95 | 3.94 | 3.96 |  | 3.95 | 3.9426 | 3.9574 |  | 3.95 | 3.9419 | 3.9581 |  | -0.0011159 | 0.0008369 | 2.8571e-06 | -8.0952e-06 | 2.381e-06 | -8.4656e-07 | 8.8183e-06 | -6.8783e-06 | 1.2346e-06 | 1.221e-07 | -8.14e-06 | 1.1752e-05 | -4.9858e-06 | 6.4103e-07