## ReadMe
INSTRUCTIONS FOR GAS DILUTION WORKBOOK
The general workflow for this workbook is from right to left. The steps involved are:
To be done annually:
Zero Gas Annual Calibration
Pollutant Annual Calibration
Annual Gas Check
To be done monthly:
Monthly Flow Check
To be done on day of use:
Day of Use
Organization
The workbook consists of several worksheets, which are displayed as tabs at the bottom of
the screen. The functions of these worksheets are described below:
Read Me(this sheet) - describes the workbook, explaining how to use the worksheets
Day of Use - to be used on day of GDS use to calculate GDS settings
Monthly Flow Check - to verify that GDS calibration has not shifted since annual calibration
Annual Gas Check - to be performed afgter annual calibration to verify that the GDS
actually produces the concentrations of diluted mixtures for both
nonreactive and reactive gases that agree with theoretical concentrations
For both the Pollutant Annual Calibration and Zero Gas Annual Calibration:
Calibration - allows for user input of calibration and other analytical data
Curves 1 - displays the calibration data, the best-fit line, and expanded uncertainties.
Residuals 1 - displays the difference between the observed responses and those
estimated by the best-fit calibration line.
Curves 2 - displays the calibration data, the best-fit quadratic curve, and
expanded uncertainties.
Residuals 2 - displays the difference between the observed responses and those
estimated by the best-fit quadratic regression line.
Chart Data - includes the data used to create the Curves and Residuals charts
Pollutant and Zero Gas Annual
Calibrations Background
This worksheet supports linear and quadratic regression models:
| Linear: | y = b0 + b1x+ e
| Quadratic: | y = b0 + b1x + b2x2 + e
The inputs are:
|  | x | the valve setting for the corresponding gas
|  | y | the measured flow rate
|  | e | a random effect due to variation of influence quantities
It is assumed that the valve setting standards (x) have negligible uncertainties.
The output quantities are estimated parameters (bs) and related uncertainties.
The workbook then helps the user perform the following functions:
- | determine which model (linear or quadratic) is best.
Conventions
The Dilution and Zero Gas Calibration sheets includes instructions that guide the user through the
steps in its use. Each worksheet is also color coded to simplify use. Shaded cells that are
bordered in blue text are for input of data. These cells are unprotected, but all other cells of
the Calibration worksheet are protected.
Derived values and statements are colored red. These cells contain formulas and are
protected to prevent alteration.
Spreadsheet background colors indicate the order of the polynomial supported by the
calculations in the area.
Light green is used for the linear model.
Yellow is used for the quadratic model.
Reference: EPA Traceability Protocol Assay and Certificatoin of Gaseous
Calibration Standards, USEPA, 2012

## Day of Use
This worksheet should only be used after filling out the Monthly Flow Check worksheet
It is only valid within one month of |  |  |  | 2012-04-30
You can only fill in the cells in:
| STEP 1 |  | Enter Pollutant Gas Concentration and Expanded Uncertainty
Enter pollutant gas concentration
Pollutant Gas Concentration
50
Enter the relative expanded uncertainty (U/conc) from the reference standard's certificate of analysis.
This value is the coverage factor (k=2) times the combined standard uncertainty (uc)
Enter the value as a decimal value (e.g., 1 percent should be entered as 0.01)
Relative Expanded Uncertainty for Pollutant Gas Concentration
0.01
| STEP 2 |  | Enter Desired Flow Rates for Pollutant and Zero Gas
Enter the desired flow rates for both the pollutant and zero gas to achieve your target dilution
and obtain the proper flow controller settings to achieve these flow rates.
|  |  | Pollutant Gas Flow Rate |  | Zero Gas Flow Rate
|  |  | 60 |  | 5940
|  |  | Pollutant Gas Flow Controller Setting |  | Zero Gas Flow Controller Setting
|  |  | 59.7678 |  | 5906.9966
| STEP 2 |  | Resulting Diluted Concentration of Pollutant and Expanded Relative Uncertainty
|  | Diluted Pollutant Gas Concentration |  |  |  | Relative Expanded Uncertainty for Diluted Conc
|  |  | 0.5 |  | 0.018599
| Relative std uncertainty of PG flow rate reference std |  |  |  | Relative std uncertainty of ZG flow rate reference std
| 0.005 |  |  |  | 0.005
| Polutant flow rate setting |  |  |  | Zero gas flow rate setting
| 59.7678 |  |  |  | 5906.9966
| Pollutant flow rate relative std uncertainty from calibration |  |  |  | Zero gas flow rate relative std uncertainty from calibration
| 0.0015157 |  |  |  | 0.0030298
| Pollutant gas flow rate relative combined std uncertainty |  |  |  | Zero gas flow rate relative combined std uncertainty
| 0.0052247 |  |  |  | 0.0058463
| Diluted pollutant gas concentration |  |  |  | Relative combined std uncertainty for diluted conc
| 0.5 |  |  |  | 0.0092993

## Monthly Flow Check
This worksheet should only be used after filling out the Annual Gas Calibration Check tab
It is only valid within one year of |  |  |  | 2012-04-29
You can only fill in the cells in:
| Date of Monthly Flow Check |  |  |  | 2012-04-30
| STEP 1 |  | Enter Flow Rate Uncertainties of Flow Rate Reference Standard
Enter expanded uncertainties (U) for both pollutant gas and zero gas flow rate reference standards
|  |  | Pollutant Gas |  | Zero Gas
|  |  | 0.01 |  | 0.01
| STEP 2 |  | Enter Flow Controller Settings
Enter flow controller setting within range of the annual calibration for both Pollutant and Zero Gas
|  |  | Pollutant Gas |  | Zero Gas
|  |  | 60 |  | 5940
|  |  | Predicted pollutant flowrate |  | Predicted zero gas flowrate
|  |  | 60.2342 |  | 5973.0165
| STEP 3 |  | Measure Flow Rate
Enter the flow rates measured by the flow rate reference standard
|  |  | Pollutant Gas |  | Zero Gas
|  |  | 60 |  | 5940
|  |  | Pollutant Passed |  | Zero Passed
|  | Predicted pollutant gas flowrate |  |  |  | Predicted zero gas flowrate
|  | 60.2342 |  |  |  | 5973.0165
|  | PG flow rate ref std relative std uncertainty |  |  |  | ZG flow rate ref std relative std uncertainty
|  | 0.005 |  |  |  | 0.005
|  | PG flow rate relative std uncertainty from calibration |  |  |  | ZG flow rate relative std uncertainty from calibration
|  | 0.001504 |  |  |  | 0.0029963
|  | PG flow rate relative combined std uncertainty |  |  |  | ZG flow rate relative combined std uncertainty
|  | 0.0052213 |  |  |  | 0.005829
|  | Pollutant flow rate acceptance criterion |  |  |  | Zero gas flow rate acceptance criterion
|  | 1.2313 |  |  |  | 129.3639

## Annual Gas Check
This worksheet should be used for the annual nonreactive gas calibratoin check (Sec. 4.1.6.3) and the annual
reactive gas calibration check (Sec. 4.1.6.4). It should only be used after filling out the Pollutant Annual
Calibration worksheet and the Zero Gas Calibration Worksheet.
You can only fill in the cells in:
| STEP 1 |  | Concentration and Uncertainty of Reference Standard
|  | Date of annual calibrations of pollutant and zero gas flow controllers | [5 empty] | 2012-04-29
|  | Fill in the certified concentration of the reference standard | [5 empty] | 50
Enter the expanded uncertainty (U) from the reference standard's certificate of analysis.
This value is the coverage factor (k=2) times the combined standard uncertainty (uc)
For example, absolute values are given in ppm and relative values are given in percent
Enter the value as a decimal value (e.g., 1 percent should be entered as 0.01)
|  | Expanded uncertainty of concentration of the reference standard | [5 empty] | 0.01
|  | Is the uncertainty absolute? (yes or no) | [5 empty] | no
It is necessary to answer whether or not this value is absolute or relative.
| STEP 2 |  | Concentration and Uncertainty of Check Standard (Lower Concentration)
|  | Enter the certified concentration of the check standard | [5 empty] | 0.5
The check standard concentration must be at least a factor of ten less than the reference
standard concentration and must fall in the well-characterized region of the callibration curve
Enter the expanded uncertainty (U) from the reference standard's certificate of analysis.
This value is the coverage factor (k=2) times the combined standard uncertainty (uc)
It is necessary to answer whether or not this value is absolute or relative.
For example, absolute values are given in ppm and relative values are given in percent
Enter the value as a decimal value (e.g., 1 percent should be entered as 0.01)
|  | Expanded uncertainty of concentration of the check standard | [5 empty] | 0.01
|  | Is the uncertainty absolute? (yes or no) | [5 empty] | no
| STEP 3 |  | Enter Settings for Pollutant and Zero Gas Flow Controllers
Select setttings for the gas flow controllers that fall in the same range as were used in
the annual calibrations and that produce a diluted concentration of the reference standard
that is equal to the undiluted concentration of the check standard.
|  | Enter the pollutant flow controller setting | [5 empty] | 60
|  | Enter the zero gas flow controller setting | [5 empty] | 5940
Predicted Concentration of Diluted
|  | Reference Standard | [5 empty] | Concentration of Check Standard
|  | 0.49919 |  |  | Match |  |  | 0.5
[5 empty] | -0.1629 | percent
Adjust the pollutant and zero gas settings so that these two concentrations match within 1 percent
Do not procede until the Match value is less than 1.0 percent
| STEP 4 |  | Measure Diluted Reference Standard and Undiluted Check Standard
Using a previously calibrated analyzer, measure the resulting diluted reference standard and
undiluted check standard. You must take at least 3 measurements of each gas mixture.
|  | Responses for Diluted Reference Standard | [5 empty] | Responses for Undiluted Check Standard
|  | 1.005 | [5 empty] | 1.005
|  | 1 | [5 empty] | 1
|  | 0.995 | [5 empty] | 0.995
|  | Mean Response to Diluted Ref Std | [5 empty] | Mean Response to Undiluted Check Std
|  | 1 |  |  | Relative Agreement between Mean Responses |  |  | 1
[5 empty] | 0 | percent
Acceptable Relative Agreement
|  | Predicted pollutant flow rate |  |  |  |  | Predicted zero gas flow rate
|  | 60.2342 |  |  |  |  | 5973.0165
|  | Relative std uncertainty for pollutant gas flow rate |  |  |  |  | Relative std uncertainty for zero gas flow rate
|  | 0.001504 |  |  |  |  | 0.0029963
|  | Reference standard concentration |  |  |  |  | Check standard concentration
|  | 50 |  |  |  |  | 0.5
|  | Predicted diluted ref std concentration |  |  |  |  | Relative concentration difference
|  | 0.49919 |  |  |  |  | -0.1629%
|  | Relative std uncertainty of ref std concentration |  |  |  |  | Relative std uncertainty of check std concentration
|  | 0.01 |  |  |  |  | 0.01
|  | Mean response for diluted reference standard |  |  |  |  | Mean response for undiluted check standard
|  | 1 |  |  |  |  | 1
|  | Standard uncertainty of ref std responses |  |  |  |  | Standard uncertainty of check std responses
|  | 0.005 |  |  |  |  | 0.005
Relative combined std uncertainty for std concentrations, flow rates, and analyzer responses
0.016163
Margin of error for confidence
0.033955
|  |  |  | Lower limit for diluted conc |  | Upper limit for diluted conc
|  |  |  | 0.96605 |  | 1.034
[5 empty] | 1 | 1 indicates diluted conc is in range, 0 indicated it is not

## Pollutant Annual Calibration
This worksheet is for the annual pollutant flow rate calibration (see Section 4.1.6.1)
| STEP 1 |  | Enter Calibration Data |  |  |  | Calculations - Linear Model |  |  |  | Calculations - Quadratic Model | [5 empty] | Calculations - Cubic Model | [5 empty] | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model |  | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model
Enter the annual pollutant flow rate calibration data in the shaded spaced below.
The first column (i) simply counts the calibration points. The second column (X)
is for the pollutant flow controller settings. The third column (Y) is for the flow
rate reference standard readings corresponding to these settings. The number of
points cannot exceed 50.
| i | Xi | Yi |  | Color Code |  | X'X |  |  |  | X2 | X'X |  |  |  |  | X3 | X'X | [7 empty] | X4 |  | X'X
| 1 | 78 | 78.4529 |  |  |  | 5 | 298 |  |  | 6084 | 5 | 298 | 18684 |  |  | 474552 | 5 | 298 | 18684 | 1222552 |  |  |  |  | 37015056 |  | 5 | 298 | 18684 | 1222552 | 82795056
| 2 | 70 | 70.1984 |  | red = derived value (protected) |  | 298 | 18684 |  |  | 4900 | 298 | 18684 | 1222552 |  |  | 343000 | 298 | 18684 | 1222552 | 82795056 |  |  |  |  | 24010000 |  | 298 | 18684 | 1222552 | 82795056 | 5760374368
| 3 | 60 | 60.2776 | [7 empty] | 3600 | 18684 | 1222552 | 82795056 |  |  | 216000 | 18684 | 1222552 | 82795056 | 5760374368 |  |  |  |  | 12960000 |  | 18684 | 1222552 | 82795056 | 5760374368 | 409225600704
| 4 | 50 | 50.1985 |  | blue = entered value (unprotected) |  | Det(X'X) |  |  |  | 2500 | [5 empty] | 125000 | 1222552 | 82795056 | 5760374368 | 409225600704 |  |  |  |  | 6250000 |  | 1222552 | 82795056 | 5760374368 | 409225600704 | 29545448854912
| 5 | 40 | 40.0262 |  |  |  | 4616 |  |  |  | 1600 | Det(X'X) |  |  |  |  | 64000 | Det(X'X) | [7 empty] | 2560000 |  | 82795056 | 5760374368 | 409225600704 | 29545448854912 | 2160172170683136
[5 empty] | black = fixed text (protected) | [6 empty] | 511072000 | [5 empty] | 5362513920023795 | [8 empty] | 1 | 2 | 3 | 4 | 5
[7 empty] | X'Xinverse | [19 empty] | 2 | X'Xinverse
[7 empty] | 4.0477 | -0.064558 |  |  |  | X'Xinverse | [5 empty] | X'Xinverse | [8 empty] | 3 | 192959.1323 | -13790.1205 | 361.5715 | -4.1279 | 0.01734
[7 empty] | -0.064558 | 0.0010832 |  |  |  | 102.3524 | -3.5822 | 0.029797 |  |  |  | 3421.0596 | -183.4245 | 3.1738 | -0.017784 | [5 empty] | 4 | -13790.1205 | 986.68 | -25.9004 | 0.29603 | -0.0012448
[12 empty] | -3.5822 | 0.12696 | -0.0010662 |  |  |  | -183.4245 | 9.8727 | -0.17144 | 0.00096375 | [5 empty] | 5 | 361.5715 | -25.9004 | 0.68068 | -0.0077887 | 3.2788e-05
[7 empty] | X'Y |  |  |  |  | 0.029797 | -0.0010662 | 9.032e-06 |  |  |  | 3.1738 | -0.17144 | 0.0029875 | -1.6848e-05 | [6 empty] | -4.1279 | 0.29603 | -0.0077887 | 8.9224e-05 | -3.7602e-07
[7 empty] | 299.1536 | [10 empty] | -0.017784 | 0.00096375 | -1.6848e-05 | 9.5305e-08 | [6 empty] | 0.01734 | -0.0012448 | 3.2788e-05 | -3.7602e-07 | 1.5863e-09
[7 empty] | 18760.8426 |  |  |  |  | X'Y | [5 empty] | X'Y
[12 empty] | 299.1536 | [5 empty] | 299.1536 | [9 empty] | X'Y |  | b
[7 empty] | b |  | b' |  |  | 18760.8426 | [5 empty] | 18760.8426 | [9 empty] | 299.1536 |  | 25.7903
[7 empty] | -0.29141 |  | -0.29141 | 1.0088 |  | 1227817.047 | [5 empty] | 1227817.047 | [9 empty] | 18760.8426 |  | -1.0022
[7 empty] | 1.0088 | [10 empty] | 83164474.6495 | [9 empty] | 1227817.047 |  | 0.056744
[12 empty] | b |  | b' |  |  |  | b | [9 empty] | 83164474.6495 |  | -0.00069398
[7 empty] | Y'Y |  |  |  |  | -0.30688 |  | -0.30688 | 1.0093 | -4.6877e-06 |  | -8.1509 | [9 empty] | 5786806804.6342 |  | 3.1051e-06
[7 empty] | 18838.0467 |  |  |  |  | 1.0093 | [5 empty] | 1.4344
[12 empty] | -4.6877e-06 | [5 empty] | -0.0074357
[7 empty] | df | [10 empty] | 4.2035e-05
[7 empty] | 3 |  |  |  |  | t(0.95,df) | [5 empty] | t(0.95,df) | [9 empty] | t(0.95,df)
[12 empty] | 4.3027 | [5 empty] | 12.7062 | [9 empty] | #NUM!
t(0.95,df)
[7 empty] | 3.1824 |  |  |  |  | b'X'Y = SSmodel | [5 empty] | b'X'Y = SSmodel | [9 empty] | b'X'Y = SSmodel
[12 empty] | 18838.0221 | [5 empty] | 18838.0407 | [9 empty] | 18838.0463
b'X'Y = SSmodel
[7 empty] | 18838.0221 |  |  |  |  | SSregression | [5 empty] | SSregression | [9 empty] | SSregression
[12 empty] | 939.4465 | [5 empty] | 939.465 | [9 empty] | 939.4707
[7 empty] | SSregression | [12 empty] | b' | [9 empty] | b'
[7 empty] | 939.4465 |  |  |  |  | SSresidual | [5 empty] | SSresidual |  | -8.1509 | 1.4344 | -0.0074357 | 4.2035e-05 |  |  |  |  | SSresidual |  | 25.7903 | -1.0022 | 0.056744 | -0.00069398 | 3.1051e-06
[12 empty] | 0.024618 | [5 empty] | 0.0060774 | [9 empty] | 0.00043908
SSresidual
[7 empty] | 0.02462 |  |  |  |  | s2 | [5 empty] | s2 | [9 empty] | s2
[12 empty] | 0.012309 | [5 empty] | 0.0060774 | [9 empty] | #DIV/0!
s2
[7 empty] | 0.0082067 |  |  |  |  | Fcrit(0.05) | [5 empty] | Fcrit(0.05) | [9 empty] | Fcrit(0.05)
[12 empty] | 19.1643 | [5 empty] | 199.5 | [9 empty] | #NUM!
Var-Covar
[7 empty] | 0.033218 | -0.00052981 |  |  |  | F | [5 empty] | F | [9 empty] | F
[7 empty] | -0.00052981 | 8.8894e-06 |  |  |  | 0.66673 | [5 empty] | 2.0253 | [9 empty] | #DIV/0!
[7 empty] | n |  |  |  |  | Preferred Model | [5 empty] | Preferred Model | [9 empty] | Preferred Model
[7 empty] | 5 |  |  |  |  | L | [5 empty] | Quadratic | [9 empty] | #DIV/0!
[12 empty] | Var-Covar | [5 empty] | Var-Covar | [9 empty] | Var-Covar
[12 empty] | 1.2598 | -0.044093 | 0.00036677 |  |  |  | 42.1093 | -2.2577 | 0.039065 | -0.00021891 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[12 empty] | -0.044093 | 0.0015627 | -1.3124e-05 |  |  |  | -2.2577 | 0.12152 | -0.0021102 | 1.1863e-05 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[12 empty] | 0.00036677 | -1.3124e-05 | 1.1117e-07 |  |  |  | 0.039065 | -0.0021102 | 3.6772e-05 | -2.0738e-07 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | -0.00021891 | 1.1863e-05 | -2.0738e-07 | 1.1731e-09 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| STEP 2 |  | Review the Parameter Estimates | [24 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Review the estimates of the parameters (b0, b1, …)
for the linear and quadratic models, their expanded
uncertainties and the overall uncertainty of the fit (u2).
The approximate level of confidence
| Linear Model |  | Expanded Uncertainty |  |  | associated with the expanded
| parameter | estimate | lower | upper |  | uncertainties is 95%.
| b0 = | -0.29141 | -0.87144 | 0.28861
| b1 = | 1.0088 | 0.99927 | 1.0182 |  | The corresponding intervals take the form
| s2 = | 0.0082067 | 0.0026336 | 0.11409 |  | estimate +/- t*u, where:
| s = | 0.090591 | 0.051319 | 0.33777 |  | estimate is the quantity of interest,
| df = | 3 |  |  |  | u is its uncertainty, and
| t = | 3.1824 |  |  |  | t is a coverage factor.
| Quadratic Model |  | Expanded Uncertainty |  |  | For each model,
| parameter | estimate | lower | upper |  | the coverage factor, t, is
| b0 = | -0.30688 | -5.7275 | 5.1138 |  | determined from the t-distribution
| b1 = | 1.0093 | 1.0026 | 1.016 |  | with appropriate degrees of
| b2 = | -4.6877e-06 | -5.1661e-06 | -4.2094e-06 |  | freedom (df).
| s2 = | 0.012309 | 0.0033002 | 0.48617
| s = | 0.11095 | 0.057448 | 0.69726 |  | Note that the uncertainties are not
| df = | 2 |  |  |  | displayed, but can easily be
| t = | 4.3027 |  |  |  | derived from the estimate, expanded
uncertainty, and the coverage factor.
Comparing the two models:
| F ratio = | 0.66673
| F critical = | 19.1643 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The linear model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then review the following estimates
for those models. Otherwise, go to Step 3.
| Cubic Model |  | Expanded Uncertainty |  |  | See above notes on calculation
| parameter | estimate | lower | upper |  | of expanded uncertainties.
| b0 = | -8.1509 | -90.6036 | 74.3018
| b1 = | 1.4344 | -2.995 | 5.8638
| b2 = | -0.0074357 | -0.084486 | 0.069615
| b3 = | 4.2035e-05 | -0.00039316 | 0.00047723
| s2 = | 0.0060774 | 0.0011929 | 6.1884
| s = | 0.077958 | 0.034539 | 2.4877
| df = | 1
| t = | 12.7062
Comparing quadratic and cubic models:
| F ratio = | 2.0253
| F critical = | 199.5 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The quadratic model appears to be a better choice than cubic.
| Quartic Model |  | Expanded Uncertainty
| parameter | estimate | lower | upper
| b0 = | 25.7903 | #NUM! | #NUM!
| b1 = | -1.0022 | #NUM! | #NUM!
| b2 = | 0.056744 | #NUM! | #NUM!
| b3 = | -0.00069398 | #NUM! | #NUM!
| b4 = | 3.1051e-06 | #NUM! | #NUM!
| s2 = | #DIV/0! | #DIV/0! | #DIV/0!
| s = | #DIV/0! | #DIV/0! | #DIV/0!
| df = | 0
| t = | #NUM!
Comparing quadratic and cubic models:
| F ratio = | #DIV/0!
| F critical = | #NUM! | (5% significance level)
#DIV/0!
#DIV/0!
| STEP 3 |  | Review the Charts
View the charts named Curves 1 and Residuals 1. Curves 1 shows bands
illustrating expanded uncertainties for the estimated regression. Compare these
bands with those of the quadratic regression, Curves 2.
(Note: You can change the associated approximate confidence level by
changing the "p-value" in cell F4 of the worksheet named Chart Data.)
Residuals 1 shows how the calibration points deviated from the calibration
line. Look for a simple pattern (such as a quadratic curve) in the chart. If
such a pattern appears, the quadratic model may be better. View Residual 2,
the deviations from the best- fit quadratic curve. If Residual 2 effectively
removes the simple pattern observed in Residual 1 and if the magnitude of
the deviations has been significantly reduced (as evidenced by a reduction
in the uncertainty u2), then the quadratic model is superior. An F-test can be
run to determine if the two overall uncertainties of fit are significantly different.
| F = | 0.66673 |  | Prob. of greater F = | 0.64642
The overall uncertainties of fit are not significantly different at the 5% level.
The linear model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then view their Curves and Residuals charts.
Note on U53: This compares the residuals of the linear and quadratic regressions. If significant, then quadratic is the preferred model.
Note on AD53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on BG53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on U56: Comparing the Linear and Quadratic solutions. Q --> quadratic is better L --> linear is better model
Note on AD56: Comparing the Quadratic and Cubic solutions
Note on BG56: Comparing the Cubic and Quartic solutions

## Zero Gas Annual Calibration
This worksheet is for the annual pollutant flow rate calibration (see Section 4.1.6.1)
| STEP 1 |  | Enter Calibration Data |  |  |  | Calculations - Linear Model |  |  |  | Calculations - Quadratic Model | [5 empty] | Calculations - Cubic Model | [5 empty] | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Cubic Model | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model |  | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model | Calculations - Quartic Model
Enter the annual zero gas flow rate calibration data in the shaded spaced below.
The first column (i) simply counts the calibration points. The second column (X)
is for the zero gas flow controller settings. The third column (Y) is for the flow
rate reference standard readings corresponding to these settings. The number of
points cannot exceed 50.
| i | Xi | Yi |  | Color Code |  | X'X |  |  |  | X2 | X'X |  |  |  |  | X3 | X'X | [7 empty] | X4 |  | X'X
| 1 | 16000 | 16023.0214 |  |  |  | 5 | 46000 |  |  | 256000000 | 5 | 46000 | 516000000 |  |  | 4096000000000 | 5 | 46000 | 516000000 | 6616000000000 |  |  |  |  | 65536000000000000 |  | 5 | 46000 | 516000000 | 6616000000000 | 91920000000000000
| 2 | 12000 | 12049.8537 |  | red = derived value (protected) |  | 46000 | 516000000 |  |  | 144000000 | 46000 | 516000000 | 6616000000000 |  |  | 1728000000000 | 46000 | 516000000 | 6616000000000 | 91920000000000000 |  |  |  |  | 20736000000000000 |  | 46000 | 516000000 | 6616000000000 | 91920000000000000 | 1.339e+21
| 3 | 8000 | 8052.8168 | [7 empty] | 64000000 | 516000000 | 6616000000000 | 91920000000000000 |  |  | 512000000000 | 516000000 | 6616000000000 | 91920000000000000 | 1.339e+21 |  |  |  |  | 4096000000000000 |  | 516000000 | 6616000000000 | 91920000000000000 | 1.339e+21 | 2.0076e+25
| 4 | 6000 | 6021.323 |  | blue = entered value (unprotected) |  | Det(X'X) |  |  |  | 36000000 | [5 empty] | 216000000000 | 6616000000000 | 91920000000000000 | 1.339e+21 | 2.0076e+25 |  |  |  |  | 1296000000000000 |  | 6616000000000 | 91920000000000000 | 1.339e+21 | 2.0076e+25 | 3.0666e+29
| 5 | 4000 | 4024.5459 |  |  |  | 464000000 |  |  |  | 16000000 | Det(X'X) |  |  |  |  | 64000000000 | Det(X'X) | [7 empty] | 256000000000000 |  | 91920000000000000 | 1.339e+21 | 2.0076e+25 | 3.0666e+29 | 4.7435e+33
[5 empty] | black = fixed text (protected) | [6 empty] | 4.8026e+23 | [5 empty] | 5.0206e+45 | [8 empty] | 1 | 2 | 3 | 4 | 5
[7 empty] | X'Xinverse | [19 empty] | 2 | X'Xinverse
[7 empty] | 1.1121 | -9.9138e-05 |  |  |  | X'Xinverse | [5 empty] | X'Xinverse | [8 empty] | 3 | 1139.72 | -0.59648 | 0.00010658 | -7.7766e-09 | 1.9797e-13
[7 empty] | -9.9138e-05 | 1.0776e-08 |  |  |  | 7.6194 | -0.0016959 | 7.9291e-08 |  |  |  | 53.4342 | -0.019299 | 2.0535e-06 | -6.6201e-11 | [5 empty] | 4 | -0.59648 | 0.00031384 | -5.6313e-08 | 4.1222e-12 | -1.0519e-16
[12 empty] | -0.0016959 | 4.0259e-07 | -1.9456e-11 |  |  |  | -0.019299 | 7.1664e-06 | -7.7799e-10 | 2.5436e-14 | [5 empty] | 5 | 0.00010658 | -5.6313e-08 | 1.0143e-11 | -7.4474e-16 | 1.9048e-20
[7 empty] | X'Y |  |  |  |  | 7.9291e-08 | -1.9456e-11 | 9.6615e-16 |  |  |  | 2.0535e-06 | -7.7799e-10 | 8.6033e-14 | -2.8526e-18 | [6 empty] | -7.7766e-09 | 4.1222e-12 | -7.4474e-16 | 5.4823e-20 | -1.4052e-24
[7 empty] | 46171.5609 | [10 empty] | -6.6201e-11 | 2.5436e-14 | -2.8526e-18 | 9.5657e-23 | [6 empty] | 1.9797e-13 | -1.0519e-16 | 1.9048e-20 | -1.4052e-24 | 3.6079e-29
[7 empty] | 517615243.2578 |  |  |  |  | X'Y | [5 empty] | X'Y
[12 empty] | 46171.5609 | [5 empty] | 46171.5609 | [9 empty] | X'Y |  | b
[7 empty] | b |  | b' |  |  | 517615243.2578 | [5 empty] | 517615243.2578 | [9 empty] | 46171.5609 |  | 513.7814
[7 empty] | 30.6556 |  | 30.6556 | 1.0004 |  | 6.6336e+12 | [5 empty] | 6.6336e+12 | [9 empty] | 517615243.2578 |  | 0.73252
[7 empty] | 1.0004 | [10 empty] | 92133661840329344 | [9 empty] | 6.6336e+12 |  | 4.925e-05
[12 empty] | b |  | b' |  |  |  | b | [9 empty] | 92133661840329344 |  | -3.6013e-09
[7 empty] | Y'Y |  |  |  |  | -35.535 |  | -35.535 | 1.0166 | -8.0652e-07 |  | 17.1282 | [9 empty] | 1.3418e+21 |  | 9.0512e-14
[7 empty] | 519237348.6787 |  |  |  |  | 1.0166 | [5 empty] | 0.9964
[12 empty] | -8.0652e-07 | [5 empty] | 1.4627e-06
[7 empty] | df | [10 empty] | -7.6096e-11
[7 empty] | 3 |  |  |  |  | t(0.95,df) | [5 empty] | t(0.95,df) | [9 empty] | t(0.95,df)
[12 empty] | 4.3027 | [5 empty] | 12.7062 | [9 empty] | #NUM!
t(0.95,df)
[7 empty] | 3.1824 |  |  |  |  | b'X'Y = SSmodel | [5 empty] | b'X'Y = SSmodel | [9 empty] | b'X'Y = SSmodel
[12 empty] | 519237061.072 | [5 empty] | 519237121.6074 | [9 empty] | 519237348.6636
b'X'Y = SSmodel
[7 empty] | 519236387.8003 |  |  |  |  | SSregression | [5 empty] | SSregression | [9 empty] | SSregression
[12 empty] | 92874454.691 | [5 empty] | 92874515.2265 | [9 empty] | 92874742.2826
[7 empty] | SSregression | [12 empty] | b' | [9 empty] | b'
[7 empty] | 92873781.4194 |  |  |  |  | SSresidual | [5 empty] | SSresidual |  | 17.1282 | 0.9964 | 1.4627e-06 | -7.6096e-11 |  |  |  |  | SSresidual |  | 513.7814 | 0.73252 | 4.925e-05 | -3.6013e-09 | 9.0512e-14
[12 empty] | 287.6067 | [5 empty] | 227.0713 | [9 empty] | 0.015125
SSresidual
[7 empty] | 960.8784 |  |  |  |  | s2 | [5 empty] | s2 | [9 empty] | s2
[12 empty] | 143.8034 | [5 empty] | 227.0713 | [9 empty] | #DIV/0!
s2
[7 empty] | 320.2928 |  |  |  |  | Fcrit(0.05) | [5 empty] | Fcrit(0.05) | [9 empty] | Fcrit(0.05)
[12 empty] | 19.1643 | [5 empty] | 199.5 | [9 empty] | #NUM!
Var-Covar
[7 empty] | 356.1877 | -0.031753 |  |  |  | F | [5 empty] | F | [9 empty] | F
[7 empty] | -0.031753 | 3.4514e-06 |  |  |  | 2.2273 | [5 empty] | 0.6333 | [9 empty] | #DIV/0!
[7 empty] | n |  |  |  |  | Preferred Model | [5 empty] | Preferred Model | [9 empty] | Preferred Model
[7 empty] | 5 |  |  |  |  | L | [5 empty] | Quadratic | [9 empty] | #DIV/0!
[12 empty] | Var-Covar | [5 empty] | Var-Covar | [9 empty] | Var-Covar
[12 empty] | 1095.6958 | -0.24388 | 1.1402e-05 |  |  |  | 7684.0193 | -2.7753 | 0.00029529 | -9.5199e-09 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[12 empty] | -0.24388 | 5.7893e-05 | -2.7979e-09 |  |  |  | -2.7753 | 0.0010305 | -1.1188e-07 | 3.6578e-12 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[12 empty] | 1.1402e-05 | -2.7979e-09 | 1.3894e-13 |  |  |  | 0.00029529 | -1.1188e-07 | 1.2372e-11 | -4.1021e-16 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[18 empty] | -9.5199e-09 | 3.6578e-12 | -4.1021e-16 | 1.3756e-20 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| STEP 2 |  | Review the Parameter Estimates | [24 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Review the estimates of the parameters (b0, b1, …)
for the linear and quadratic models, their expanded
uncertainties and the overall uncertainty of the fit (u2).
The approximate level of confidence
| Linear Model |  | Expanded Uncertainty |  |  | associated with the expanded
| parameter | estimate | lower | upper |  | uncertainties is 95%.
| b0 = | 30.6556 | -29.4065 | 90.7177
| b1 = | 1.0004 | 0.99449 | 1.0063 |  | The corresponding intervals take the form
| s2 = | 320.2928 | 102.7853 | 4452.7312 |  | estimate +/- t*u, where:
| s = | 17.8967 | 10.1383 | 66.7288 |  | estimate is the quantity of interest,
| df = | 3 |  |  |  | u is its uncertainty, and
| t = | 3.1824 |  |  |  | t is a coverage factor.
| Quadratic Model |  | Expanded Uncertainty |  |  | For each model,
| parameter | estimate | lower | upper |  | the coverage factor, t, is
| b0 = | -35.535 | -4749.9335 | 4678.8635 |  | determined from the t-distribution
| b1 = | 1.0166 | 1.0164 | 1.0169 |  | with appropriate degrees of
| b2 = | -8.0652e-07 | -8.0653e-07 | -8.0652e-07 |  | freedom (df).
| s2 = | 143.8034 | 38.5563 | 5679.9295
| s = | 11.9918 | 6.2094 | 75.3653 |  | Note that the uncertainties are not
| df = | 2 |  |  |  | displayed, but can easily be
| t = | 4.3027 |  |  |  | derived from the estimate, expanded
uncertainty, and the coverage factor.
Comparing the two models:
| F ratio = | 2.2273
| F critical = | 19.1643 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The linear model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then review the following estimates
for those models. Otherwise, go to Step 3.
| Cubic Model |  | Expanded Uncertainty |  |  | See above notes on calculation
| parameter | estimate | lower | upper |  | of expanded uncertainties.
| b0 = | 17.1282 | -1096.6792 | 1130.9355
| b1 = | 0.9964 | 0.58851 | 1.4043
| b2 = | 1.4627e-06 | -4.323e-05 | 4.6155e-05
| b3 = | -7.6096e-11 | -1.5663e-09 | 1.4142e-09
| s2 = | 227.0713 | 44.5707 | 231217.2055
| s = | 15.0689 | 6.6761 | 480.8505
| df = | 1
| t = | 12.7062
Comparing quadratic and cubic models:
| F ratio = | 0.6333
| F critical = | 199.5 | (5% significance level)
The overall uncertainties of fit are not significantly different at the 5% level.
The quadratic model appears to be a better choice than cubic.
| Quartic Model |  | Expanded Uncertainty
| parameter | estimate | lower | upper
| b0 = | 513.7814 | #NUM! | #NUM!
| b1 = | 0.73252 | #NUM! | #NUM!
| b2 = | 4.925e-05 | #NUM! | #NUM!
| b3 = | -3.6013e-09 | #NUM! | #NUM!
| b4 = | 9.0512e-14 | #NUM! | #NUM!
| s2 = | #DIV/0! | #DIV/0! | #DIV/0!
| s = | #DIV/0! | #DIV/0! | #DIV/0!
| df = | 0
| t = | #NUM!
Comparing quadratic and cubic models:
| F ratio = | #DIV/0!
| F critical = | #NUM! | (5% significance level)
#DIV/0!
#DIV/0!
| STEP 3 |  | Review the Charts
View the charts named Curves 1 and Residuals 1. Curves 1 shows bands
illustrating expanded uncertainties for the estimated regression. Compare these
bands with those of the quadratic regression, Curves 2.
(Note: You can change the associated approximate confidence level by
changing the "p-value" in cell F4 of the worksheet named Chart Data.)
Residuals 1 shows how the calibration points deviated from the calibration
line. Look for a simple pattern (such as a quadratic curve) in the chart. If
such a pattern appears, the quadratic model may be better. View Residual 2,
the deviations from the best- fit quadratic curve. If Residual 2 effectively
removes the simple pattern observed in Residual 1 and if the magnitude of
the deviations has been significantly reduced (as evidenced by a reduction
in the uncertainty u2), then the quadratic model is superior. An F-test can be
run to determine if the two overall uncertainties of fit are significantly different.
| F = | 2.2273 |  | Prob. of greater F = | 0.32481
The overall uncertainties of fit are not significantly different at the 5% level.
The linear model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then view their Curves and Residuals charts.
Note on U53: This compares the residuals of the linear and quadratic regressions. If significant, then quadratic is the preferred model.
Note on AD53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on BG53: This compares the residuals of the quadratic and cubic regressions. If significant, then cubic is the preferred model.
Note on U56: Comparing the Linear and Quadratic solutions. Q --> quadratic is better L --> linear is better model
Note on AD56: Comparing the Quadratic and Cubic solutions
Note on BG56: Comparing the Cubic and Quartic solutions

## Chart Data Pollutant
This sheet contains the chart data. The chart data include initial calibration data and the
regression information derived from those data.
Only the initial calibration data (Step 1 in Measurement Data worksheet) are included.
| Recalibration data (Step 6) are not included. | [5 empty] | 0.95 | = Confidence Level for Plotting Confidence Bands
40 | = minimum concentration | [5 empty] | 3.1824 | = t for linear model
78 | = maximum concentration | [5 empty] | 4.3027 | = t for quadratic model
5 | = n, the number of initial calibration points | [5 empty] | 12.7062 | = t for cubic model
[7 empty] | #NUM! | = t for quartic model
Initial Calibration Data |  |  |  |  | Estimates, Linear Model |  |  |  | Estimates, Quadratic Model |  |  |  | Estimates, Cubic Model |  |  |  | Estimates, Quartic Model
i | Xi | Yi |  |  | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual
1 | 78 | 78.4529 | [5 empty] | 0.060988 |  |  |  | 0.061809 |  |  |  | 0.012775 |  |  |  | 2.2708e-06
2 | 70 | 70.1984 | [5 empty] | -0.12348 |  |  |  | -0.12378 |  |  |  | -0.040779 |  |  |  | 1.758e-06
3 | 60 | 60.2776 | [5 empty] | 0.043391 |  |  |  | 0.042528 |  |  |  | 0.054372 |  |  |  | 1.2361e-06
4 | 50 | 50.1985 | [5 empty] | 0.051906 |  |  |  | 0.05142 |  |  |  | -0.034954 |  |  |  | 8.3264e-07
5 | 40 | 40.0262 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
6 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
7 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
8 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
9 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
10 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
11 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
12 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
13 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
14 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
15 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
16 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
17 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
18 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
19 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
20 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
21 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
22 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
23 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
24 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
25 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
26 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
27 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
28 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
29 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
30 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
31 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
32 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
33 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
34 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
35 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
36 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
37 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
38 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
39 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
40 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
41 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
42 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
43 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
44 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
45 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
46 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
47 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
48 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
49 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
50 | -999 | -999 | [5 empty] | -0.032804 |  |  |  | -0.031974 |  |  |  | 0.0085851 |  |  |  | 5.2998e-07
[21 empty] | x'V for Linear |  | x'V for Quadratic |  |  | x'V for Cubic |  |  |  | x'V for Quartic
1 | 40 | 1600 | 64000 | 2560000 | 40.059 | 39.8327 | 40.2853 |  | 40.0582 | 39.6055 | 40.5108 |  | 40.0176 | 39.6246 | 40.4107 |  | 40.0262 | #DIV/0! | #DIV/0! |  | 0.012026 | -0.00017423 | 0.082962 | -0.0025846 | 1.968e-05 | 0.29418 | -0.014031 | 0.00021978 | -1.1319e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 40.76 | 1661.3776 | 67717.751 | 2760175.5298 | 40.8257 | 40.6053 | 41.0461 |  | 40.825 | 40.4023 | 41.2477 |  | 40.8076 | 40.4539 | 41.1614 |  | 40.7888 | #DIV/0! | #DIV/0! |  | 0.011623 | -0.00016748 | 0.071963 | -0.0022025 | 1.6529e-05 | 0.1622 | -0.0070925 | 0.00010201 | -4.8357e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 41.52 | 1723.9104 | 71576.7598 | 2971867.0672 | 41.5923 | 41.3777 | 41.8069 |  | 41.5918 | 41.196 | 41.9875 |  | 41.595 | 41.267 | 41.923 |  | 41.5548 | #DIV/0! | #DIV/0! |  | 0.01122 | -0.00016072 | 0.061388 | -0.0018355 | 1.3507e-05 | 0.044422 | -0.00091615 | -2.5659e-06 | 9.0917e-08 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 42.28 | 1787.5984 | 75579.6604 | 3195508.0397 | 42.359 | 42.1501 | 42.5679 |  | 42.3585 | 41.9866 | 42.7305 |  | 42.3798 | 42.0655 | 42.6942 |  | 42.3235 | #DIV/0! | #DIV/0! |  | 0.010818 | -0.00015396 | 0.051237 | -0.0014838 | 1.0612e-05 | -0.059726 | 0.0045294 | -9.4508e-05 | 5.9463e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 43.04 | 1852.4416 | 79729.0865 | 3431539.8814 | 43.1256 | 42.9224 | 43.3289 |  | 43.1253 | 42.7739 | 43.4768 |  | 43.1622 | 42.852 | 43.4725 |  | 43.0945 | #DIV/0! | #DIV/0! |  | 0.010415 | -0.00014721 | 0.041509 | -0.0011471 | 7.8469e-06 | -0.15082 | 0.0092754 | -0.00017436 | 1.0307e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 43.8 | 1918.44 | 84027.672 | 3680412.0336 | 43.8923 | 43.6946 | 44.09 |  | 43.8921 | 43.5578 | 44.2263 |  | 43.9423 | 43.6293 | 44.2553 |  | 43.8675 | #DIV/0! | #DIV/0! |  | 0.010012 | -0.00014045 | 0.032205 | -0.00082568 | 5.2098e-06 | -0.22944 | 0.013353 | -0.00024266 | 1.4021e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 44.56 | 1985.5936 | 88478.0508 | 3942581.9444 | 44.659 | 44.4666 | 44.8513 |  | 44.6588 | 44.3385 | 44.9792 |  | 44.7202 | 44.4005 | 45.0399 |  | 44.6419 | #DIV/0! | #DIV/0! |  | 0.0096097 | -0.0001337 | 0.023324 | -0.00051939 | 2.701e-06 | -0.29616 | 0.016794 | -0.00029996 | 1.7121e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 45.32 | 2053.9024 | 93082.8568 | 4218515.0687 | 45.4256 | 45.2386 | 45.6127 |  | 45.4256 | 45.1159 | 45.7353 |  | 45.496 | 45.1679 | 45.824 |  | 45.4175 | #DIV/0! | #DIV/0! |  | 0.009207 | -0.00012694 | 0.014867 | -0.00022825 | 3.2072e-07 | -0.35156 | 0.019628 | -0.00034681 | 1.9636e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 46.08 | 2123.3664 | 97844.7237 | 4508684.8686 | 46.1923 | 46.0104 | 46.3742 |  | 46.1924 | 45.8902 | 46.4945 |  | 46.2698 | 45.9333 | 46.6062 |  | 46.1938 | #DIV/0! | #DIV/0! |  | 0.0088044 | -0.00012018 | 0.0068342 | 4.7717e-05 | -1.9312e-06 | -0.39621 | 0.021889 | -0.00038375 | 2.1598e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 46.84 | 2193.9856 | 102766.2855 | 4813572.813 | 46.9589 | 46.7821 | 47.1358 |  | 46.9591 | 46.6616 | 47.2566 |  | 47.0417 | 46.6978 | 47.3855 |  | 46.9706 | #DIV/0! | #DIV/0! |  | 0.0084017 | -0.00011343 | -0.00077522 | 0.00030853 | -4.0546e-06 | -0.43069 | 0.023606 | -0.00041133 | 2.3038e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 47.6 | 2265.76 | 107850.176 | 5133668.3776 | 47.7256 | 47.5536 | 47.8976 |  | 47.7258 | 47.4305 | 48.0212 |  | 47.8118 | 47.4623 | 48.1613 |  | 47.7475 | #DIV/0! | #DIV/0! |  | 0.0079991 | -0.00010667 | -0.0079609 | 0.00055418 | -6.0496e-06 | -0.45557 | 0.02481 | -0.00043009 | 2.3987e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 48.36 | 2338.6896 | 113099.0291 | 5469469.0451 | 48.4923 | 48.3249 | 48.6596 |  | 48.4926 | 48.1973 | 48.7879 |  | 48.5803 | 48.2271 | 48.9335 |  | 48.5242 | #DIV/0! | #DIV/0! |  | 0.0075964 | -9.9917e-05 | -0.014723 | 0.00078466 | -7.9162e-06 | -0.47144 | 0.025534 | -0.00044059 | 2.4475e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 49.12 | 2412.7744 | 118515.4785 | 5821480.3053 | 49.2589 | 49.0961 | 49.4217 |  | 49.2593 | 48.9623 | 49.5563 |  | 49.3472 | 48.9926 | 49.7018 |  | 49.3005 | #DIV/0! | #DIV/0! |  | 0.0071937 | -9.3161e-05 | -0.021061 | 0.00099999 | -9.6544e-06 | -0.47887 | 0.025809 | -0.00044336 | 2.4533e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 49.88 | 2488.0144 | 124102.1583 | 6190215.6546 | 50.0256 | 49.867 | 50.1841 |  | 50.026 | 49.726 | 50.3261 |  | 50.1127 | 49.7589 | 50.4665 |  | 50.0761 | #DIV/0! | #DIV/0! |  | 0.0067911 | -8.6405e-05 | -0.026976 | 0.0012002 | -1.1264e-05 | -0.47844 | 0.025665 | -0.00043895 | 2.4193e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 50.64 | 2564.4096 | 129861.7021 | 6576196.5966 | 50.7922 | 50.6378 | 50.9467 |  | 50.7928 | 50.4887 | 51.0968 |  | 50.8769 | 50.5259 | 51.2279 |  | 50.8509 | #DIV/0! | #DIV/0! |  | 0.0063884 | -7.9649e-05 | -0.032467 | 0.0013852 | -1.2745e-05 | -0.47071 | 0.025134 | -0.00042792 | 2.3485e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 51.4 | 2641.96 | 135796.744 | 6979952.6416 | 51.5589 | 51.4083 | 51.7095 |  | 51.5595 | 51.2508 | 51.8681 |  | 51.6399 | 51.2936 | 51.9862 |  | 51.6246 | #DIV/0! | #DIV/0! |  | 0.0059858 | -7.2893e-05 | -0.037534 | 0.001555 | -1.4098e-05 | -0.45628 | 0.024247 | -0.0004108 | 2.244e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 52.16 | 2720.6656 | 141909.9177 | 7402021.307 | 52.3255 | 52.1785 | 52.4725 |  | 52.3262 | 52.0126 | 52.6398 |  | 52.4018 | 52.0617 | 52.7418 |  | 52.3971 | #DIV/0! | #DIV/0! |  | 0.0055831 | -6.6137e-05 | -0.042177 | 0.0017097 | -1.5323e-05 | -0.43571 | 0.023036 | -0.00038814 | 2.1089e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 52.92 | 2800.5264 | 148203.8571 | 7842948.1171 | 53.0922 | 52.9485 | 53.2359 |  | 53.0929 | 52.7743 | 53.4114 |  | 53.1626 | 52.8299 | 53.4953 |  | 53.1683 | #DIV/0! | #DIV/0! |  | 0.0051805 | -5.9381e-05 | -0.046397 | 0.0018492 | -1.6419e-05 | -0.40959 | 0.021531 | -0.00036049 | 1.9463e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 53.68 | 2881.5424 | 154681.196 | 8303286.603 | 53.8589 | 53.7182 | 53.9995 |  | 53.8596 | 53.5362 | 54.183 |  | 53.9226 | 53.5981 | 54.2471 |  | 53.938 | #DIV/0! | #DIV/0! |  | 0.0047778 | -5.2625e-05 | -0.050193 | 0.0019736 | -1.7386e-05 | -0.37848 | 0.019764 | -0.00032839 | 1.7592e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 54.44 | 2963.7136 | 161344.5684 | 8783598.3028 | 54.6255 | 54.4876 | 54.7634 |  | 54.6263 | 54.2985 | 54.9541 |  | 54.6819 | 54.3659 | 54.9978 |  | 54.7062 | #DIV/0! | #DIV/0! |  | 0.0043752 | -4.5869e-05 | -0.053566 | 0.0020828 | -1.8226e-05 | -0.34297 | 0.017766 | -0.00029239 | 1.5509e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 55.2 | 3047.04 | 168196.608 | 9284452.7616 | 55.3922 | 55.2567 | 55.5277 |  | 55.393 | 55.0612 | 55.7247 |  | 55.4404 | 55.133 | 55.7479 |  | 55.4728 | #DIV/0! | #DIV/0! |  | 0.0039725 | -3.9113e-05 | -0.056515 | 0.0021768 | -1.8936e-05 | -0.30363 | 0.015568 | -0.00025304 | 1.3243e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 55.96 | 3131.5216 | 175239.9487 | 9806427.5313 | 56.1588 | 56.0254 | 56.2923 |  | 56.1597 | 55.8246 | 56.4947 |  | 56.1984 | 55.8988 | 56.498 |  | 56.2378 | #DIV/0! | #DIV/0! |  | 0.0035698 | -3.2357e-05 | -0.05904 | 0.0022557 | -1.9519e-05 | -0.26104 | 0.013202 | -0.00021089 | 1.0825e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 56.72 | 3217.1584 | 182477.2244 | 10350108.1707 | 56.9255 | 56.7937 | 57.0573 |  | 56.9263 | 56.5886 | 57.264 |  | 56.956 | 56.6633 | 57.2488 |  | 57.0011 | #DIV/0! | #DIV/0! |  | 0.0031672 | -2.5602e-05 | -0.061141 | 0.0023194 | -1.9973e-05 | -0.21578 | 0.010699 | -0.00016647 | 8.2868e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 57.48 | 3303.9504 | 189911.069 | 10916088.2457 | 57.6921 | 57.5617 | 57.8226 |  | 57.693 | 57.3535 | 58.0326 |  | 57.7133 | 57.4259 | 58.0007 |  | 57.7629 | #DIV/0! | #DIV/0! |  | 0.0027645 | -1.8846e-05 | -0.062819 | 0.002368 | -2.0298e-05 | -0.16842 | 0.0080904 | -0.00012034 | 5.6589e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 58.24 | 3391.8976 | 197544.1162 | 11504969.3289 | 58.4588 | 58.3292 | 58.5884 |  | 58.4597 | 58.1191 | 58.8002 |  | 58.4703 | 58.1865 | 58.7541 |  | 58.523 | #DIV/0! | #DIV/0! |  | 0.0023619 | -1.209e-05 | -0.064073 | 0.0024013 | -2.0495e-05 | -0.11953 | 0.0054068 | -7.3035e-05 | 2.972e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 59 | 3481 | 205379 | 12117361 | 59.2255 | 59.0964 | 59.3545 |  | 59.2263 | 58.8856 | 59.567 |  | 59.2273 | 58.945 | 59.5095 |  | 59.2816 | #DIV/0! | #DIV/0! |  | 0.0019592 | -5.3336e-06 | -0.064903 | 0.0024196 | -2.0564e-05 | -0.069704 | 0.0026797 | -2.5112e-05 | 2.5728e-08 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 59.76 | 3571.2576 | 213418.3542 | 12753880.8456 | 59.9921 | 59.8632 | 60.1211 |  | 59.993 | 59.653 | 60.333 |  | 59.9842 | 59.7014 | 60.267 |  | 60.0388 | #DIV/0! | #DIV/0! |  | 0.0015566 | 1.4223e-06 | -0.065309 | 0.0024226 | -2.0504e-05 | -0.019506 | -5.9437e-05 | 2.2888e-05 | -2.4545e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 60.52 | 3662.6704 | 221664.8126 | 13415154.459 | 60.7588 | 60.6296 | 60.888 |  | 60.7596 | 60.4212 | 61.098 |  | 60.7412 | 60.4559 | 61.0266 |  | 60.7947 | #DIV/0! | #DIV/0! |  | 0.0011539 | 8.1783e-06 | -0.065292 | 0.0024106 | -2.0316e-05 | 0.030483 | -0.0027795 | 7.0417e-05 | -5.1325e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 61.28 | 3755.2384 | 230121.0092 | 14101815.4408 | 61.5254 | 61.3955 | 61.6554 |  | 61.5263 | 61.1903 | 61.8622 |  | 61.4985 | 61.2088 | 61.7883 |  | 61.5494 | #DIV/0! | #DIV/0! |  | 0.00075126 | 1.4934e-05 | -0.064851 | 0.0023833 | -1.9999e-05 | 0.079687 | -0.0054493 | 0.00011693 | -7.7456e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 62.04 | 3848.9616 | 238789.5777 | 14814505.3983 | 62.2921 | 62.1611 | 62.4231 |  | 62.2929 | 61.9602 | 62.6257 |  | 62.2561 | 61.9605 | 62.5518 |  | 62.303 | #DIV/0! | #DIV/0! |  | 0.00034861 | 2.169e-05 | -0.063987 | 0.0023409 | -1.9554e-05 | 0.12753 | -0.0080375 | 0.00016188 | -1.0263e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 62.8 | 3943.84 | 247673.152 | 15553873.9456 | 63.0588 | 62.9263 | 63.1912 |  | 63.0596 | 62.7308 | 63.3883 |  | 63.0142 | 62.7117 | 63.3168 |  | 63.0559 | #DIV/0! | #DIV/0! |  | -5.4048e-05 | 2.8446e-05 | -0.062699 | 0.0022833 | -1.898e-05 | 0.17344 | -0.010513 | 0.00020472 | -1.2654e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 63.56 | 4039.8736 | 256774.366 | 16320578.704 | 63.8254 | 63.6911 | 63.9597 |  | 63.8262 | 63.502 | 64.1503 |  | 63.7728 | 63.4628 | 64.0829 |  | 63.8082 | #DIV/0! | #DIV/0! |  | -0.0004567 | 3.5202e-05 | -0.060987 | 0.0022106 | -1.8278e-05 | 0.21683 | -0.012844 | 0.00024491 | -1.4888e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 64.32 | 4137.0624 | 266095.8536 | 17115285.3015 | 64.5921 | 64.4556 | 64.7286 |  | 64.5928 | 64.2738 | 64.9118 |  | 64.5321 | 64.2145 | 64.8498 |  | 64.5601 | #DIV/0! | #DIV/0! |  | -0.00085936 | 4.1958e-05 | -0.058851 | 0.0021227 | -1.7448e-05 | 0.25713 | -0.015 | 0.00028189 | -1.6933e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 65.08 | 4235.4064 | 275640.2485 | 17938667.3732 | 65.3587 | 65.2197 | 65.4978 |  | 65.3594 | 65.046 | 65.6729 |  | 65.2922 | 64.9673 | 65.6171 |  | 65.3119 | #DIV/0! | #DIV/0! |  | -0.001262 | 4.8714e-05 | -0.056292 | 0.0020196 | -1.6489e-05 | 0.29376 | -0.01695 | 0.00031513 | -1.8759e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 65.84 | 4334.9056 | 285410.1847 | 18791406.5609 | 66.1254 | 65.9835 | 66.2673 |  | 66.126 | 65.8183 | 66.4337 |  | 66.0532 | 65.7218 | 66.3845 |  | 66.0639 | #DIV/0! | #DIV/0! |  | -0.0016647 | 5.547e-05 | -0.053309 | 0.0019014 | -1.5402e-05 | 0.32615 | -0.018662 | 0.00034408 | -2.0335e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 66.6 | 4435.56 | 295408.296 | 19674192.5136 | 66.892 | 66.747 | 67.0371 |  | 66.8926 | 66.5907 | 67.1946 |  | 66.8151 | 66.4786 | 67.1517 |  | 66.8163 | #DIV/0! | #DIV/0! |  | -0.0020673 | 6.2226e-05 | -0.049902 | 0.001768 | -1.4186e-05 | 0.35372 | -0.020104 | 0.00036818 | -2.1629e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 67.36 | 4537.3696 | 305637.2163 | 20587722.887 | 67.6587 | 67.5102 | 67.8072 |  | 67.6592 | 67.3628 | 67.9557 |  | 67.5782 | 67.2381 | 67.9184 |  | 67.5697 | #DIV/0! | #DIV/0! |  | -0.00247 | 6.8982e-05 | -0.046072 | 0.0016195 | -1.2842e-05 | 0.37589 | -0.021247 | 0.0003869 | -2.2612e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 68.12 | 4640.3344 | 316099.5793 | 21532703.3438 | 68.4254 | 68.2732 | 68.5775 |  | 68.4258 | 68.1343 | 68.7174 |  | 68.3425 | 68.0006 | 68.6844 |  | 68.3242 | #DIV/0! | #DIV/0! |  | -0.0028726 | 7.5738e-05 | -0.041818 | 0.0014558 | -1.137e-05 | 0.39208 | -0.022058 | 0.00039969 | -2.3252e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 68.88 | 4744.4544 | 326798.0191 | 22509847.5537 | 69.192 | 69.0359 | 69.3482 |  | 69.1924 | 68.9048 | 69.48 |  | 69.1082 | 68.7666 | 69.4498 |  | 69.0803 | #DIV/0! | #DIV/0! |  | -0.0032753 | 8.2494e-05 | -0.03714 | 0.0012769 | -9.7686e-06 | 0.40173 | -0.022506 | 0.00040599 | -2.3518e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 69.64 | 4849.7296 | 337735.1693 | 23519877.1931 | 69.9587 | 69.7984 | 70.119 |  | 69.959 | 69.674 | 70.244 |  | 69.8752 | 69.5362 | 70.2143 |  | 69.8384 | #DIV/0! | #DIV/0! |  | -0.0036779 | 8.925e-05 | -0.032039 | 0.0010829 | -8.0392e-06 | 0.40425 | -0.02256 | 0.00040528 | -2.338e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 70.4 | 4956.16 | 348913.664 | 24563521.9456 | 70.7253 | 70.5606 | 70.89 |  | 70.7256 | 70.4415 | 71.0097 |  | 70.6439 | 70.3096 | 70.9782 |  | 70.599 | #DIV/0! | #DIV/0! |  | -0.0040806 | 9.6006e-05 | -0.026513 | 0.00087372 | -6.1814e-06 | 0.39907 | -0.022189 | 0.00039699 | -2.2806e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 71.16 | 5063.7456 | 360336.1369 | 25641519.5015 | 71.492 | 71.3227 | 71.6613 |  | 71.4922 | 71.2067 | 71.7776 |  | 71.4142 | 71.0867 | 71.7417 |  | 71.3624 | #DIV/0! | #DIV/0! |  | -0.0044833 | 0.00010276 | -0.020565 | 0.00064938 | -4.1952e-06 | 0.3856 | -0.021361 | 0.00038059 | -2.1766e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 71.92 | 5172.4864 | 372005.2219 | 26754615.5582 | 72.2587 | 72.0846 | 72.4327 |  | 72.2587 | 71.9693 | 72.5482 |  | 72.1863 | 71.8672 | 72.5053 |  | 72.1293 | #DIV/0! | #DIV/0! |  | -0.0048859 | 0.00010952 | -0.014192 | 0.00040987 | -2.0805e-06 | 0.36328 | -0.020046 | 0.00035552 | -2.0228e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 72.68 | 5282.3824 | 383923.5528 | 27903563.8198 | 73.0253 | 72.8463 | 73.2043 |  | 73.0253 | 72.729 | 73.3216 |  | 72.9602 | 72.6504 | 73.2701 |  | 72.9 | #DIV/0! | #DIV/0! |  | -0.0052886 | 0.00011627 | -0.0073957 | 0.0001552 | 1.6257e-07 | 0.33153 | -0.018211 | 0.00032124 | -1.8162e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 73.44 | 5393.4336 | 396093.7636 | 29089125.9976 | 73.792 | 73.6079 | 73.976 |  | 73.7919 | 73.4854 | 74.0983 |  | 73.7362 | 73.4352 | 74.0372 |  | 73.6752 | #DIV/0! | #DIV/0! |  | -0.0056912 | 0.00012303 | -0.00017577 | -0.00011463 | 2.5341e-06 | 0.28976 | -0.015827 | 0.00027721 | -1.5537e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 74.2 | 5505.64 | 408518.488 | 30312071.8096 | 74.5586 | 74.3694 | 74.7479 |  | 74.5584 | 74.2384 | 74.8784 |  | 74.5143 | 74.2198 | 74.8088 |  | 74.4554 | #DIV/0! | #DIV/0! |  | -0.0060939 | 0.00012979 | 0.0074679 | -0.00039962 | 5.034e-06 | 0.23741 | -0.012861 | 0.00022287 | -1.2323e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 74.96 | 5619.0016 | 421200.3599 | 31573178.9808 | 75.3253 | 75.1307 | 75.5199 |  | 75.325 | 74.9879 | 75.662 |  | 75.2946 | 75.0017 | 75.5874 |  | 75.2412 | #DIV/0! | #DIV/0! |  | -0.0064965 | 0.00013654 | 0.015535 | -0.00069977 | 7.6624e-06 | 0.1739 | -0.0092817 | 0.00015769 | -8.4866e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 75.72 | 5733.5184 | 434142.0132 | 32873233.2431 | 76.0919 | 75.8919 | 76.292 |  | 76.0915 | 75.734 | 76.449 |  | 76.0772 | 75.7781 | 76.3762 |  | 76.0332 | #DIV/0! | #DIV/0! |  | -0.0068992 | 0.0001433 | 0.024026 | -0.0010151 | 1.0419e-05 | 0.098649 | -0.0050589 | 8.1113e-05 | -3.9989e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 76.48 | 5849.1904 | 447346.0818 | 34213028.3355 | 76.8586 | 76.653 | 77.0642 |  | 76.858 | 76.4768 | 77.2393 |  | 76.8622 | 76.5461 | 77.1784 |  | 76.8321 | #DIV/0! | #DIV/0! |  | -0.0073018 | 0.00015005 | 0.032941 | -0.0013456 | 1.3304e-05 | 0.011081 | -0.00016094 | -7.4049e-06 | 1.1715e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 77.24 | 5966.0176 | 460815.1994 | 35593366.0035 | 77.6253 | 77.414 | 77.8365 |  | 77.6246 | 77.2163 | 78.0329 |  | 77.6499 | 77.3037 | 77.996 |  | 77.6384 | #DIV/0! | #DIV/0! |  | -0.0077045 | 0.00015681 | 0.042279 | -0.0016912 | 1.6318e-05 | -0.08938 | 0.0054435 | -0.00010841 | 7.0555e-07 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 78 | 6084 | 474552 | 37015056 | 78.3919 | 78.1749 | 78.609 |  | 78.3911 | 77.9527 | 78.8295 |  | 78.4401 | 78.05 | 78.8303 |  | 78.4529 | #DIV/0! | #DIV/0! |  | -0.0081071 | 0.00016357 | 0.052042 | -0.002052 | 1.946e-05 | -0.20331 | 0.011786 | -0.00022245 | 1.3684e-06 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## Chart Data Zero Gas
This sheet contains the chart data. The chart data include initial calibration data and the
regression information derived from those data.
Only the initial calibration data (Step 1 in Measurement Data worksheet) are included.
| Recalibration data (Step 6) are not included. | [5 empty] | 0.95 | = Confidence Level for Plotting Confidence Bands
4000 | = minimum concentration | [5 empty] | 3.1824 | = t for linear model
16000 | = maximum concentration | [5 empty] | 4.3027 | = t for quadratic model
5 | = n, the number of initial calibration points | [5 empty] | 12.7062 | = t for cubic model
[7 empty] | #NUM! | = t for quartic model
Initial Calibration Data |  |  |  |  | Estimates, Linear Model |  |  |  | Estimates, Quadratic Model |  |  |  | Estimates, Cubic Model |  |  |  | Estimates, Quartic Model
i | Xi | Yi |  |  | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual
1 | 16000 | 16023.0214 | [5 empty] | -13.9934 |  |  |  | -1.2003 |  |  |  | 0.65332 |  |  |  | -3.8908e-07
2 | 12000 | 12049.8537 | [5 empty] | 14.4286 |  |  |  | 1.858 |  |  |  | -3.2666 |  |  |  | 6.7374e-07
3 | 8000 | 8052.8168 | [5 empty] | 18.9816 |  |  |  | 6.8559 |  |  |  | 9.7998 |  |  |  | 8.4332e-07
4 | 6000 | 6021.323 | [5 empty] | -11.7173 |  |  |  | -13.9422 |  |  |  | -10.4531 |  |  |  | 7.2244e-07
5 | 4000 | 4024.5459 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
6 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
7 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
8 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
9 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
10 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
11 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
12 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
13 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
14 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
15 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
16 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
17 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
18 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
19 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
20 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
21 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
22 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
23 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
24 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
25 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
26 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
27 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
28 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
29 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
30 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
31 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
32 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
33 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
34 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
35 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
36 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
37 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
38 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
39 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
40 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
41 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
42 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
43 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
44 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
45 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
46 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
47 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
48 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
49 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
50 | -999 | -999 | [5 empty] | -7.6995 |  |  |  | 6.4286 |  |  |  | 3.2666 |  |  |  | 5.2181e-07
[21 empty] | x'V for Linear |  | x'V for Quadratic |  |  | x'V for Cubic |  |  |  | x'V for Quartic
1 | 4000 | 16000000 | 64000000000 | 256000000000000 | 4032.2454 | 3992.3206 | 4072.1703 |  | 4018.1173 | 3972.3199 | 4063.9148 |  | 4021.2793 | 3979.5501 | 4063.0086 |  | 4024.5459 | #DIV/0! | #DIV/0! |  | 229.175 | -0.017947 | 302.631 | -0.057069 | 2.4338e-06 | 698.2032 | -0.20906 | 1.9479e-05 | -5.7159e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 4240 | 17977600 | 76225024000 | 323194101760000 | 4272.3408 | 4233.4981 | 4311.1835 |  | 4260.5158 | 4218.1942 | 4302.8373 |  | 4262.3788 | 4225.8536 | 4298.904 |  | 4259.8035 | #DIV/0! | #DIV/0! |  | 221.5543 | -0.017119 | 266.6501 | -0.048708 | 2.037e-06 | 499.7206 | -0.13826 | 1.208e-05 | -3.3678e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 4480 | 20070400 | 89915392000 | 402820956160000 | 4512.4362 | 4474.6533 | 4550.2191 |  | 4502.8213 | 4463.6154 | 4542.0272 |  | 4503.5353 | 4470.8187 | 4536.2518 |  | 4496.583 | #DIV/0! | #DIV/0! |  | 213.9335 | -0.016291 | 231.9827 | -0.040669 | 1.6563e-06 | 321.3059 | -0.07499 | 5.5053e-06 | -1.2907e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 4720 | 22278400 | 105154048000 | 496327106560000 | 4752.5316 | 4715.7843 | 4789.2789 |  | 4745.0339 | 4708.5579 | 4781.5099 |  | 4744.7424 | 4714.4812 | 4775.0037 |  | 4734.7164 | #DIV/0! | #DIV/0! |  | 206.3127 | -0.015462 | 198.6289 | -0.032952 | 1.2916e-06 | 162.1695 | -0.018943 | -2.7945e-07 | 5.2682e-11 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 4960 | 24601600 | 122023936000 | 605238722560000 | 4992.627 | 4956.8889 | 5028.365 |  | 4987.1536 | 4952.9975 | 5021.3096 |  | 4985.994 | 4956.9888 | 5014.9992 |  | 4974.0434 | #DIV/0! | #DIV/0! |  | 198.692 | -0.014634 | 166.5887 | -0.025558 | 9.4288e-07 | 21.5218 | 0.030181 | -5.3081e-06 | 2.0962e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 5200 | 27040000 | 140608000000 | 731161600000000 | 5232.7224 | 5197.9649 | 5267.4798 |  | 5229.1803 | 5196.915 | 5261.4457 |  | 5227.2836 | 5198.5936 | 5255.9735 |  | 5214.4104 | #DIV/0! | #DIV/0! |  | 191.0712 | -0.013806 | 135.862 | -0.018486 | 6.1017e-07 | -101.4268 | 0.072688 | -9.6146e-06 | 3.4287e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 5440 | 29593600 | 160989184000 | 875781160960000 | 5472.8178 | 5439.0098 | 5506.6257 |  | 5471.1142 | 5440.3004 | 5501.928 |  | 5468.6049 | 5439.5924 | 5497.6175 |  | 5455.6713 | #DIV/0! | #DIV/0! |  | 183.4505 | -0.012977 | 106.4488 | -0.011736 | 2.9346e-07 | -207.4659 | 0.10888 | -1.3233e-05 | 4.5359e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 5680 | 32262400 | 183250432000 | 1040862453760000 | 5712.9131 | 5680.0209 | 5745.8054 |  | 5712.9552 | 5683.1575 | 5742.7528 |  | 5709.9518 | 5680.2587 | 5739.6448 |  | 5697.6873 | #DIV/0! | #DIV/0! |  | 175.8297 | -0.012149 | 78.3492 | -0.0053089 | -7.2362e-09 | -297.3852 | 0.13906 | -1.6198e-05 | 5.4292e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 5920 | 35046400 | 207474688000 | 1228250152960000 | 5953.0085 | 5920.9952 | 5985.0219 |  | 5954.7032 | 5925.5068 | 5983.8996 |  | 5951.3177 | 5920.8051 | 5981.8303 |  | 5940.3265 | #DIV/0! | #DIV/0! |  | 168.2089 | -0.011321 | 51.5632 | 0.00079616 | -2.9193e-07 | -371.9743 | 0.16353 | -1.8542e-05 | 6.12e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 6160 | 37945600 | 233744896000 | 1439868559360000 | 6193.1039 | 6161.9297 | 6224.2781 |  | 6196.3583 | 6167.3853 | 6225.3314 |  | 6192.6964 | 6161.3772 | 6224.0157 |  | 6183.4645 | #DIV/0! | #DIV/0! |  | 160.5882 | -0.010492 | 26.0907 | 0.0065789 | -5.6062e-07 | -432.0228 | 0.1826 | -2.0301e-05 | 6.6196e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 6400 | 40960000 | 262144000000 | 1677721600000000 | 6433.1993 | 6402.8211 | 6463.5776 |  | 6437.9205 | 6408.8432 | 6466.9978 |  | 6434.0817 | 6402.0643 | 6466.0991 |  | 6426.9838 | #DIV/0! | #DIV/0! |  | 152.9674 | -0.009664 | 1.9317 | 0.012039 | -8.133e-07 | -478.3203 | 0.19657 | -2.1507e-05 | 6.9395e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 6640 | 44089600 | 292754944000 | 1943892828160000 | 6673.2947 | 6643.6658 | 6702.9236 |  | 6679.3898 | 6649.9391 | 6708.8406 |  | 6675.4671 | 6642.9132 | 6708.0211 |  | 6670.7744 | #DIV/0! | #DIV/0! |  | 145.3467 | -0.0088357 | -20.9137 | 0.017177 | -1.05e-06 | -511.6564 | 0.20574 | -2.2196e-05 | 7.091e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 6880 | 47334400 | 325660672000 | 2240545423360000 | 6913.3901 | 6884.4603 | 6942.3198 |  | 6920.7662 | 6890.7332 | 6950.7993 |  | 6916.8465 | 6883.94 | 6949.753 |  | 6914.7334 | #DIV/0! | #DIV/0! |  | 137.7259 | -0.0080073 | -42.4456 | 0.021993 | -1.2707e-06 | -532.8207 | 0.21041 | -2.2401e-05 | 7.0857e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 7120 | 50694400 | 360944128000 | 2569922191360000 | 7153.4855 | 7125.2009 | 7181.77 |  | 7162.0497 | 7131.2832 | 7192.8162 |  | 7158.2134 | 7125.139 | 7191.2877 |  | 7158.765 | #DIV/0! | #DIV/0! |  | 130.1051 | -0.007179 | -62.6639 | 0.026487 | -1.4753e-06 | -542.6029 | 0.21089 | -2.2156e-05 | 6.9349e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 7360 | 54169600 | 398688256000 | 2934345564160000 | 7393.5809 | 7365.8838 | 7421.278 |  | 7403.2403 | 7371.6412 | 7434.8394 |  | 7399.5615 | 7366.4888 | 7432.6343 |  | 7402.7808 | #DIV/0! | #DIV/0! |  | 122.4844 | -0.0063506 | -81.5687 | 0.030658 | -1.664e-06 | -541.7926 | 0.20749 | -2.1495e-05 | 6.6501e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 7600 | 57760000 | 438976000000 | 3336217600000000 | 7633.6763 | 7606.5051 | 7660.8474 |  | 7644.338 | 7611.8517 | 7676.8242 |  | 7640.8846 | 7607.956 | 7673.8133 |  | 7646.6995 | #DIV/0! | #DIV/0! |  | 114.8636 | -0.0055223 | -99.1599 | 0.034507 | -1.8366e-06 | -531.1794 | 0.2005 | -2.0452e-05 | 6.2425e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 7840 | 61465600 | 481890304000 | 3778019983360000 | 7873.7716 | 7847.0614 | 7900.4819 |  | 7885.3427 | 7851.9523 | 7918.7331 |  | 7882.1764 | 7849.4987 | 7914.8541 |  | 7890.4469 | #DIV/0! | #DIV/0! |  | 107.2429 | -0.0046939 | -115.4376 | 0.038033 | -1.9933e-06 | -511.5528 | 0.19023 | -1.9062e-05 | 5.7237e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 8080 | 65286400 | 527514112000 | 4262314024960000 | 8113.867 | 8087.5491 | 8140.1849 |  | 8126.2545 | 8091.9735 | 8160.5356 |  | 8123.4305 | 8091.0682 | 8155.7927 |  | 8133.9563 | #DIV/0! | #DIV/0! |  | 99.6221 | -0.0038656 | -130.4018 | 0.041237 | -2.1339e-06 | -483.7026 | 0.17699 | -1.7358e-05 | 5.1051e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 8320 | 69222400 | 575930368000 | 4791740661760000 | 8353.9624 | 8327.9652 | 8379.9596 |  | 8367.0735 | 8331.9399 | 8402.2071 |  | 8364.6405 | 8332.6114 | 8396.6697 |  | 8377.1679 | #DIV/0! | #DIV/0! |  | 92.0013 | -0.0030373 | -144.0523 | 0.044119 | -2.2586e-06 | -448.4184 | 0.16107 | -1.5374e-05 | 4.398e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 8560 | 73273600 | 627222016000 | 5369020456960000 | 8594.0578 | 8568.3071 | 8619.8085 |  | 8607.7995 | 8571.8709 | 8643.7281 |  | 8605.8003 | 8574.0731 | 8637.5275 |  | 8620.0292 | #DIV/0! | #DIV/0! |  | 84.3806 | -0.0022089 | -156.3894 | 0.046679 | -2.3672e-06 | -406.4897 | 0.14278 | -1.3144e-05 | 3.6139e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 8800 | 77440000 | 681472000000 | 5996953600000000 | 8834.1532 | 8808.5724 | 8859.734 |  | 8848.4326 | 8811.7818 | 8885.0834 |  | 8846.9035 | 8815.399 | 8878.4079 |  | 8862.4951 | #DIV/0! | #DIV/0! |  | 76.7598 | -0.0013806 | -167.4129 | 0.048916 | -2.4598e-06 | -358.7061 | 0.12242 | -1.0703e-05 | 2.7641e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 9040 | 81721600 | 738763264000 | 6678419906560000 | 9074.2486 | 9048.7598 | 9099.7374 |  | 9088.9728 | 9051.6842 | 9126.2613 |  | 9087.9437 | 9056.5386 | 9119.3489 |  | 9104.5275 | #DIV/0! | #DIV/0! |  | 69.1391 | -0.00055223 | -177.1228 | 0.050831 | -2.5365e-06 | -305.8574 | 0.10029 | -8.0836e-06 | 1.8602e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 9280 | 86118400 | 799178752000 | 7416378818560000 | 9314.344 | 9288.8684 | 9339.8196 |  | 9329.42 | 9291.5868 | 9367.2533 |  | 9328.9147 | 9297.4487 | 9360.3808 |  | 9346.0955 | #DIV/0! | #DIV/0! |  | 61.5183 | 0.00027611 | -185.5192 | 0.052424 | -2.5971e-06 | -248.7331 | 0.076712 | -5.321e-06 | 9.1342e-11 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 9520 | 90630400 | 862801408000 | 8213869404160000 | 9554.4394 | 9528.898 | 9579.9807 |  | 9569.7744 | 9531.4959 | 9608.053 |  | 9569.8102 | 9538.0973 | 9601.5231 |  | 9587.1755 | #DIV/0! | #DIV/0! |  | 53.8975 | 0.0011045 | -192.6021 | 0.053694 | -2.6417e-06 | -188.1228 | 0.051973 | -2.4487e-06 | -6.4724e-12 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 9760 | 95257600 | 929714176000 | 9074010357760000 | 9794.5347 | 9768.8492 | 9820.2202 |  | 9810.0359 | 9771.4152 | 9848.6565 |  | 9810.6238 | 9778.4655 | 9842.7821 |  | 9827.751 | #DIV/0! | #DIV/0! |  | 46.2768 | 0.0019328 | -198.3714 | 0.054642 | -2.6703e-06 | -124.8161 | 0.02638 | 4.992e-07 | -1.0628e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 10000 | 100000000 | 1000000000000 | 10000000000000000 | 10034.6301 | 10008.7235 | 10060.5368 |  | 10050.2044 | 10011.3469 | 10089.0619 |  | 10051.3493 | 10018.5486 | 10084.1499 |  | 10067.8129 | #DIV/0! | #DIV/0! |  | 38.656 | 0.0027611 | -202.8271 | 0.055268 | -2.6829e-06 | -59.6027 | 0.00023652 | 3.4887e-06 | -2.0695e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 10240 | 104857600 | 1073741824000 | 10995116277760000 | 10274.7255 | 10248.5226 | 10300.9284 |  | 10290.28 | 10251.291 | 10329.2691 |  | 10291.9802 | 10258.3559 | 10325.6045 |  | 10307.3591 | #DIV/0! | #DIV/0! |  | 31.0353 | 0.0035895 | -205.9693 | 0.055571 | -2.6795e-06 | 6.7278 | -0.026154 | 6.4856e-06 | -3.0734e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 10480 | 109830400 | 1151022592000 | 12062716764160000 | 10514.8209 | 10488.2492 | 10541.3926 |  | 10530.2628 | 10491.246 | 10569.2795 |  | 10532.5104 | 10497.9082 | 10567.1126 |  | 10546.3949 | #DIV/0! | #DIV/0! |  | 23.4145 | 0.0044178 | -207.798 | 0.055552 | -2.6601e-06 | 73.3859 | -0.052487 | 9.4561e-06 | -4.063e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 10720 | 114918400 | 1231925248000 | 13206238658560000 | 10754.9163 | 10727.9062 | 10781.9264 |  | 10770.1526 | 10731.2084 | 10809.0967 |  | 10772.9334 | 10737.2352 | 10808.6317 |  | 10784.9326 | #DIV/0! | #DIV/0! |  | 15.7937 | 0.0052462 | -208.3131 | 0.055211 | -2.6247e-06 | 139.5818 | -0.078461 | 1.2366e-05 | -5.0269e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 10960 | 120121600 | 1316532736000 | 14429198786560000 | 10995.0117 | 10967.4969 | 11022.5265 |  | 11009.9495 | 10971.1729 | 11048.726 |  | 11013.2431 | 10976.3726 | 11050.1136 |  | 11022.9918 | #DIV/0! | #DIV/0! |  | 8.173 | 0.0060745 | -207.5147 | 0.054547 | -2.5733e-06 | 204.526 | -0.10377 | 1.5182e-05 | -5.9538e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 11200 | 125440000 | 1404928000000 | 15735193600000000 | 11235.1071 | 11207.0249 | 11263.1892 |  | 11249.6534 | 11211.1324 | 11288.1745 |  | 11253.433 | 11215.359 | 11291.5069 |  | 11260.5994 | #DIV/0! | #DIV/0! |  | 0.55223 | 0.0069029 | -205.4027 | 0.053561 | -2.5058e-06 | 267.4288 | -0.12811 | 1.7869e-05 | -6.8322e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 11440 | 130873600 | 1497193984000 | 17127899176960000 | 11475.2025 | 11446.494 | 11503.911 |  | 11489.2645 | 11451.0771 | 11527.4519 |  | 11493.4969 | 11454.2341 | 11532.7596 |  | 11497.7894 | #DIV/0! | #DIV/0! |  | -7.0685 | 0.0077312 | -201.9772 | 0.052253 | -2.4224e-06 | 327.5008 | -0.15119 | 2.0393e-05 | -7.6508e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 11680 | 136422400 | 1593413632000 | 18611071221760000 | 11715.2978 | 11685.9078 | 11744.6879 |  | 11728.7827 | 11690.9951 | 11766.5703 |  | 11733.4284 | 11693.0368 | 11773.82 |  | 11734.603 | #DIV/0! | #DIV/0! |  | -14.6893 | 0.0085595 | -197.2381 | 0.050623 | -2.323e-06 | 383.9521 | -0.17269 | 2.2721e-05 | -8.398e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 11920 | 142086400 | 1693669888000 | 20188545064960000 | 11955.3932 | 11925.2701 | 11985.5163 |  | 11968.2079 | 11930.8714 | 12005.5445 |  | 11973.2212 | 11931.8039 | 12014.6386 |  | 11971.0886 | #DIV/0! | #DIV/0! |  | -22.31 | 0.0093879 | -191.1855 | 0.04867 | -2.2075e-06 | 435.9932 | -0.19231 | 2.4818e-05 | -9.0625e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 12160 | 147865600 | 1798045696000 | 21864235663360000 | 12195.4886 | 12164.5847 | 12226.3926 |  | 12207.5403 | 12170.6879 | 12244.3926 |  | 12212.8691 | 12170.5696 | 12255.1686 |  | 12207.3019 | #DIV/0! | #DIV/0! |  | -29.9308 | 0.010216 | -183.8193 | 0.046395 | -2.0761e-06 | 482.8345 | -0.20975 | 2.665e-05 | -9.6329e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 12400 | 153760000 | 1906624000000 | 23642137600000000 | 12435.584 | 12403.855 | 12467.3131 |  | 12446.7797 | 12410.4228 | 12483.1366 |  | 12452.3657 | 12409.3648 | 12495.3666 |  | 12443.3058 | #DIV/0! | #DIV/0! |  | -37.5516 | 0.011045 | -175.1396 | 0.043797 | -1.9286e-06 | 523.6864 | -0.22471 | 2.8184e-05 | -1.0098e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 12640 | 159769600 | 2019487744000 | 25526325084160000 | 12675.6794 | 12643.0843 | 12708.2745 |  | 12685.9262 | 12650.0502 | 12721.8022 |  | 12691.7047 | 12648.2161 | 12735.1933 |  | 12679.1702 | #DIV/0! | #DIV/0! |  | -45.1723 | 0.011873 | -165.1464 | 0.040877 | -1.7652e-06 | 557.7593 | -0.23689 | 2.9385e-05 | -1.0446e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 12880 | 165894400 | 2136719872000 | 27520951951360000 | 12915.7748 | 12882.276 | 12949.2736 |  | 12924.9798 | 12889.54 | 12960.4197 |  | 12930.8798 | 12887.145 | 12974.6145 |  | 12914.9726 | #DIV/0! | #DIV/0! |  | -52.7931 | 0.012701 | -153.8396 | 0.037635 | -1.5857e-06 | 584.2635 | -0.24597 | 3.0219e-05 | -1.0665e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 13120 | 172134400 | 2258403328000 | 29630251663360000 | 13155.8702 | 13121.4328 | 13190.3075 |  | 13163.9405 | 13128.8578 | 13199.0232 |  | 13169.8846 | 13126.167 | 13213.6023 |  | 13150.7972 | #DIV/0! | #DIV/0! |  | -60.4138 | 0.01353 | -141.2192 | 0.034071 | -1.3902e-06 | 602.4094 | -0.25165 | 3.0653e-05 | -1.0745e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 13360 | 178489600 | 2384621056000 | 31858537308160000 | 13395.9656 | 13360.5577 | 13431.3734 |  | 13402.8083 | 13367.9652 | 13437.6514 |  | 13408.713 | 13365.2892 | 13452.1367 |  | 13386.7359 | #DIV/0! | #DIV/0! |  | -68.0346 | 0.014358 | -127.2853 | 0.030184 | -1.1788e-06 | 611.4075 | -0.25364 | 3.0652e-05 | -1.0674e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 13600 | 184960000 | 2515456000000 | 34210201600000000 | 13636.061 | 13599.6531 | 13672.4688 |  | 13641.5831 | 13606.8209 | 13676.3454 |  | 13647.3584 | 13604.5082 | 13690.2087 |  | 13622.8875 | #DIV/0! | #DIV/0! |  | -75.6554 | 0.015186 | -112.0378 | 0.025975 | -9.5128e-07 | 610.468 | -0.25163 | 3.0182e-05 | -1.044e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 13840 | 191545600 | 2650991104000 | 36689716879360000 | 13876.1563 | 13838.7215 | 13913.5912 |  | 13880.2651 | 13845.3821 | 13915.1481 |  | 13885.8148 | 13843.8054 | 13927.8242 |  | 13859.3581 | #DIV/0! | #DIV/0! |  | -83.2761 | 0.016015 | -95.4768 | 0.021444 | -7.0779e-07 | 598.8014 | -0.24532 | 2.9209e-05 | -1.0032e-09 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 14080 | 198246400 | 2791309312000 | 39301635112960000 | 14116.2517 | 14077.7649 | 14154.7386 |  | 14118.8541 | 14083.6065 | 14154.1018 |  | 14124.0756 | 14083.141 | 14165.0102 |  | 14096.261 | #DIV/0! | #DIV/0! |  | -90.8969 | 0.016843 | -77.6023 | 0.01659 | -4.483e-07 | 575.618 | -0.2344 | 2.7699e-05 | -9.4388e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 14320 | 205062400 | 2936493568000 | 42050587893760000 | 14356.3471 | 14316.7854 | 14395.9089 |  | 14357.3503 | 14321.4553 | 14393.2452 |  | 14362.1347 | 14322.4453 | 14401.8241 |  | 14333.7167 | #DIV/0! | #DIV/0! |  | -98.5176 | 0.017671 | -58.4142 | 0.011414 | -1.7281e-07 | 540.1283 | -0.21856 | 2.5618e-05 | -8.6487e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 14560 | 211993600 | 3086626816000 | 44941286440960000 | 14596.4425 | 14555.7847 | 14637.1003 |  | 14595.7535 | 14558.8957 | 14632.6113 |  | 14599.9857 | 14561.6056 | 14638.3657 |  | 14571.8529 | #DIV/0! | #DIV/0! |  | -106.1384 | 0.0185 | -37.9126 | 0.0059157 | 1.1869e-07 | 491.5427 | -0.19752 | 2.2933e-05 | -7.6504e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 14800 | 219040000 | 3241792000000 | 47978521600000000 | 14836.5379 | 14794.7647 | 14878.3111 |  | 14834.0638 | 14795.9025 | 14872.2251 |  | 14837.6222 | 14800.4508 | 14874.7937 |  | 14810.8046 | #DIV/0! | #DIV/0! |  | -113.7592 | 0.019328 | -16.0974 | 9.5051e-05 | 4.262e-07 | 429.0714 | -0.17095 | 1.9609e-05 | -6.4325e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 15040 | 226201600 | 3402072064000 | 51167163842560000 | 15076.6333 | 15033.7267 | 15119.5399 |  | 15072.2812 | 15032.4598 | 15112.1026 |  | 15075.0381 | 15038.7369 | 15111.3393 |  | 15050.7138 | #DIV/0! | #DIV/0! |  | -121.3799 | 0.020156 | 7.0313 | -0.0060479 | 7.4971e-07 | 351.9249 | -0.13857 | 1.5611e-05 | -4.9836e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 15280 | 233478400 | 3567549952000 | 54512163266560000 | 15316.7287 | 15272.6721 | 15360.7852 |  | 15310.4057 | 15268.5607 | 15352.2507 |  | 15312.2269 | 15276.1449 | 15348.3089 |  | 15291.73 | #DIV/0! | #DIV/0! |  | -129.0007 | 0.020985 | 31.4736 | -0.012513 | 1.0892e-06 | 259.3136 | -0.10006 | 1.0907e-05 | -3.2922e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 15520 | 240870400 | 3738308608000 | 58018549596160000 | 15556.8241 | 15511.6023 | 15602.0458 |  | 15548.4373 | 15504.2059 | 15592.6687 |  | 15549.1824 | 15512.3105 | 15586.0543 |  | 15534.0096 | #DIV/0! | #DIV/0! |  | -136.6214 | 0.021813 | 57.2294 | -0.019301 | 1.4447e-06 | 150.4479 | -0.055118 | 5.4615e-06 | -1.347e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 15760 | 248377600 | 3914430976000 | 61691432181760000 | 15796.9194 | 15750.5183 | 15843.3206 |  | 15786.3759 | 15739.4022 | 15833.3497 |  | 15785.8982 | 15746.894 | 15824.9025 |  | 15777.7164 | #DIV/0! | #DIV/0! |  | -144.2422 | 0.022641 | 84.2988 | -0.026411 | 1.8163e-06 | 24.5381 | -0.0034487 | -7.5882e-07 | 8.6352e-11 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1 | 16000 | 256000000 | 4096000000000 | 65536000000000000 | 16037.0148 | 15989.4213 | 16084.6084 |  | 16024.2217 | 15974.1607 | 16074.2827 |  | 16022.3681 | 15979.6626 | 16065.0736 |  | 16023.0214 | #DIV/0! | #DIV/0! |  | -151.863 | 0.02347 | 112.6817 | -0.033843 | 2.2038e-06 | -119.2054 | 0.055255 | -7.7882e-06 | 3.3507e-10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!